## Form Responses
Timestamp | Email Address | System Deployed by | Name of the Hospital | City | System Identity | What is the identity of other system (Exp, Customer System or System Number) | Please provide the software version | Extra instruments provided with the system (Please write NA, if not available) | Upload system photograph | Date of surgery | Application Engineer(s) | Name of the Surgeon | Type of surgery | What is the other type of surgery | Total Number of X-Rays During surgery (Approximate, if exact number is not known. Write 1000, if Not applicable, or, Unknown) | Tool Tip Accuracy in mm (Specify 100.00 if not available) | Fudicial Registration in mm (Specify 100.00, if not available) | Surface Registration in mm (Specify 100.00 if not available) | Please provide navigation Accuracy and other relevant Photographs from surgery | Issues we faced during surgery/Our observations: Please provide information in the list format about 1. Software, 2. Instruments, 3. Data, and 4. Others. Please write NA if there is no information about an item. | Feedback from surgeon (Write NA, if not given) Please arrange feedback in a list covering: 1. Accuracy, 2. Ergonomics, 3. Ease of Use, 4. Anything other - please specify. If there is no feedback about any aspect, please write NA for that aspect. | Feedback form photo | Was surgery successful? | What was exact description of the failure (What was supposed to happen and what was supposed to happen)? | In your opinion, what was the reason for failure of the surgery? | Any other information | Any loss/damage to instruments seen after the surgery? (Write NA, if not seen). Provide Details in short. | Did you face any issues while moving the system? (Write NA, if no issues were faced). | What are your inputs/requests to design and developments, logistics, assembly, etc.? (Please write NA, if you do not have any inputs at this stage) | What are the inputs/requests from the Surgeon and/or hospital staff to the company? (Write NA, if none is provided.) | Upload Surgery Photos with the Doctor | Upload Demo Request Form | Upload consent Form | Software Related Issues | Instruments Related Issues | Data Related Issues | Other Issues | Surgeon Feedback related to Accuracy | Surgeon Feedback related to Ergonomics | Surgeon Feedback related to Ease of use | Surgeon Feedback related to Anything else | Please provide summary of the procedure (Whatever you think is significant point and can help in improving system and demo practices. If you do not have anything specific to add, please respond NA). | System Base - Optical, or, EM | Comments by Dr. Kurwale ?? | Comments By ??
2024 Demo Data
2024-06-27 15:24 | janhavi1999@gmail.com | suraj | For Aluminium EM Box & New Trolley - 207 | Mumbai | NS TRL MS 01 | Mobile System |  | NA |  | 2024-06-13 | Janhavi and suraj | Dr. Majza Turel | Tumour Surgery |  |  | 0.23 | 00 | 0.23 |  | NA | NA |  | Yes |  |  |  | Na | NA | NA | NA
2024-07-06 13:12 | gurjit@rxoom.con | Gurjit Singh | For Trolley - 209 | Other (Please Specify) | NS TRL 02 MF 204 |  |  | Na | https://drive.google.com/open?id=1ZKdHpmBmY8CSRsfXwfuzWT1GNf3wBW2N | 2024-07-04 | Gurjit Singh | Dr Jain Georgon | X-Ray Spine |  | 25 | 0.2 | 100 | 0.73 | https://drive.google.com/open?id=1L4-M9tsqPKxJ1NmKB0IVhf_HSqwViehj | High Cuve on operating area cause issue to attach acyerlic plate | Fusion of ct spine and X-ray is required for bettering patient care. | https://drive.google.com/open?id=1puMdlNj8Wdqpg5WX2D3bDn1rYBM2LhZe | Yes |  |  |  | Na | Na | I think we need to design references points/something which can be easily fit on curve on patient spine. | From Dr. Ct scan and X-ray scan fusion required for better patient care.
2024-07-06 14:02 | Gurjit@rxoom.com | Gurjit Singh | MOSC Hospital | Other (Please Specify) | NS TRL 02 MF 204 |  |  | Na |  | 2024-07-05 | Gurjit Singh | Jain George | CT-Spine |  |  | 0.12 | 5.86 | 0.85 | https://drive.google.com/open?id=1Le7qEXph7pR1ilr2zZQ5CqtzF4x6enTa | Reference get undetected after some time of registration | In spine surgery with minimal exposure navigation is not useful | https://drive.google.com/open?id=1KRS_mgyHUdQ8kNyg29N4fWNtm1MnTfpg | No | Take extra incision for Ct reference/expose spine to collect fiducial points. Dr. kept spine minimal exposed and didn’t able to collect fiducial points accurately which caused higher fiducial accuracy error. | Dr want to keep the spine minimal exposed while doing the surgery using ct scan. Dr mentioned it a limitation in our spine module. |  | Na | Na | NA | Fusion of X-ray and ct.
2024-09-13 12:19 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Brains super speciality hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1UxNh8m98X4NsHmJc7ys5G2yi0vrajtCf | 2024-09-11 | Hemarshitha H S | Dr Anirudh and Dr Venkataramana | Cranial |  |  | 0.25 | 4.94 | 0.65 | https://drive.google.com/open?id=1F9e1NnCzrj-SQ1AuIBu35R4HI7rULlE9 | NA | During the registration process, Dr. Anirudh mentioned that their team had been using Stryker since 2014 and Medtronic since 2020, but found the registration process in those systems difficult especially in prone position, which led to reduced usage. He expressed his happiness with our system's ease of use and accuracy. Later, Dr. Venkataramana used our system for the surgery and gave excellent feedback, stating that our system is the best compared to others. | https://drive.google.com/open?id=1pscdFachsjBN37_XU90MFJEKe-Oq3tbO | Yes |  |  | No | NA | NA | NA | NA
2024-09-18 17:09 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Fortis hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1pvjMoZtqXH0isXQhKCEYs-2U5wiOBdC0 | 2024-09-16 | Hemarshitha H S | Dr Vikas Naik | Cranial |  |  | 0.37 | 5.51 | 0.78 | https://drive.google.com/open?id=139SjZX0IwJtkVrWjtU3xuJ3867ibL_65 | NA | Surgeon was happy with system performance |  | Yes |  |  | NA | NA | NA | NA | NA
2024-09-13 12:19 | hemarshitha@rxoom.com | Hemarshitha H. S. | Brains super speciality hospital | Bangalore | NS TRL 02 MF 204 |  | 1.02 | NA |  |  | https://drive.google.com/open?id=1UxNh8m98X4NsHmJc7ys5G2yi0vrajtCf | 2024-09-11 | Hemarshitha H S |  | Dr. Anirudh and Dr Venkataramana | Cranial |  |  | 0.25 | 4.94 | 0.65 |  | Yes |  | https://drive.google.com/open?id=1F9e1NnCzrj-SQ1AuIBu35R4HI7rULlE9 | NA | During the registration process, Dr. Anirudh mentioned that their team had been using Stryker since 2014 and Medtronic since 2020, but found the registration process in those systems difficult especially in prone position, which led to reduced usage. He expressed his happiness with our system's ease of use and accuracy. Later, Dr. Venkataramana used our system for the surgery and gave excellent feedback, stating that our system is the best compared to others. | https://drive.google.com/open?id=1pscdFachsjBN37_XU90MFJEKe-Oq3tbO | Yes | NA | NA | NA
2024-09-18 17:09 | hemarshitha@rxoom.com | Hemarshitha H. S. | Fortis hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA |  |  | https://drive.google.com/open?id=1pvjMoZtqXH0isXQhKCEYs-2U5wiOBdC0 | 2024-09-16 | Hemarshitha H S |  | Dr. Vikas Naik | Cranial |  |  | 0.37 | 5.51 | 0.78 |  | Yes |  | https://drive.google.com/open?id=139SjZX0IwJtkVrWjtU3xuJ3867ibL_65 | NA | Surgeon was happy with system performance |  | Yes | NA | NA | NA
2024-09-21 14:41 | trushal@rxoom.com | Trushal Gulhane | BMHRC | Bhopal | NS TRL 01 MF 101 |  | V1.2 | NA |  |  | https://drive.google.com/open?id=1G9B3wZu0eRPfhKCI3xw8N7Ct09Ntmczi | 2024-09-20 | Trushal Gulhane |  | Dr. Sandeep sorte | Cranial |  |  | 0.3 | 2.6 | 0.8 |  | Yes |  | https://drive.google.com/open?id=1_3YPxhtYLE--Y3fRzxI6aQs5mrxSCxMV | NA | NA | https://drive.google.com/open?id=1uuaBUA836GkcLvLRAGfbi4Adpn1-fWmz | Yes | NA | NA | NA
2024-09-23 22:40 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderbad | NS TRL 03 SE 101 |  | 2.1 | NA |  |  | https://drive.google.com/open?id=1f-3yiMleyYaTrZgqnX5nQrXhCgXJW3ZM | 2024-09-20 | Rohan Ghosh |  | Dr Venkatesh Yaddula | Cranial |  |  | 0.03 | 6.99 | 1.16 |  | Yes |  | https://drive.google.com/open?id=1F1oiFMS--bCTc-pnElFaqLU9ym0erI8Z | CSF has been leaked & tumor is size is very small as per doctor's expectation & anatomy was very critical. | Our demo system was used successfully in an endoscopic Tran nasal Tumor excision. The doctor created a model based on T2 data and found the pre-draped accuracy to be correct. However, post-draped accuracy was affected due to CSF leakage during the surgery. Despite this, the doctor was overall satisfied with the system's performance. The doctor appreciated the system's ease of use and accuracy, and believes that it is a valuable tool for neurosurgeons. | https://drive.google.com/open?id=1mY29i4Tp6ZA0ykE8OGfkOhJhbQKlL8Gx | Yes | The doctor was overall satisfied with the system's performance. The doctor appreciated the system's ease of use and accuracy, and believes that it is a valuable tool for neurosurgeons. | NA | NA
2024-09-26 22:04 | trushal@rxoom.com | Trushal Gulhane | BMHRC | Bhopal | NS TRL 01 MF 101 |  | V1.2 | Biopsy needle provided |  |  | https://drive.google.com/open?id=1cM-7Yxl5WNL1M0FgfPuq8ijfKqjZBH_v | 2024-09-26 | Trushal Gulhane |  | Dr. Sandeep sorte | Biopsy |  |  | 0.3 | 3.4 | 0.6 |  | Yes |  | https://drive.google.com/open?id=1TrZdbeoG3wsm1CNIEytfNcUzaiqrmMns | 1. Our demo system was used successfully in brain biopsy. First the doctor planned the trajectory and registered the patient, after draping the patient doctor was aligning the biopsy tracker with planned the trajectory but ball and socket joint near the biopsy tracker was not free to move and it was less degrees of rotation for placing the tracker, eventually after some trial align the biopsy tracker with trajectory and did a small bit drill through the pipe of biopsy tool. With the help of a biopsy tool give the projection to the tool till the target and put the stopper on the needle ( projection+ tool length). And also tracked the needle. 2. The doctor marked the trajectory on axial Slice however that same trajectory was visible on another/adjacent/near of axial slices where trajectory marked. | NA | https://drive.google.com/open?id=11yuL1ptpQ6zdoX5GXvdwsLpONJvDDuWx | Yes | NA | NA | NA
2024-09-27 10:28 | manoj@rxoom.com | Suraj Garad | Sahyadri hospital | Pune | NS TRL 02 MF 201 |  | 1.1 | NA |  |  | https://drive.google.com/open?id=1omTm8sblePOghGGc4LeUXDKb5aIwkszW | 2024-09-23 | Manoj dhavale ,suraj gard, anita nimbalkar |  | Dr. Ranjeet deshmukh | Cranial |  |  | 0.9 | 3 | 0.59 |  | Yes |  | https://drive.google.com/open?id=1Sb7_yXgRMeO0necfgVttJKjKOVDJRbKz | NA | good | https://drive.google.com/open?id=1FXgoWGhcax7ytH9y8lS2vFfFhTs3zvL2 | Yes | NA | NA | NA
2024-09-28 10:22 | Janhavi@rxoom.com | Janhvi Barhate | Zyloma Shalby Hospital | Mumbai | NS TRL 01 MF 101 |  | R1.2 | Biopsy Needle |  |  | https://drive.google.com/open?id=10aky-Tig2D50MEWr_Atix3gMv-JZdBS6 | 2024-09-27 | Janhavi Barhate |  | Dr. Shashank Joshi | Biopsy |  |  | 0.07 | 000 | 0.9 |  | Yes |  | https://drive.google.com/open?id=1LHnkONHCcdQnZCQzBgoHAyuFvSfaAgoR | NA | Dr. Shashank Joshi wants a Needle | https://drive.google.com/open?id=1LYXWbAfLwTSVPhtxobBr_PAcWiOVbG_L | Yes | NA | NA | NA
2024-09-28 10:55 | Gurjit@rxoom.com | Gurjit Singh | Globle Hospital | Mumbai | Mobile System |  | R 1.2 | NA |  |  | https://drive.google.com/open?id=1PRtYTdry6Pa3qbRhxe3NgbVJUQzQfTOx | 9/14/0024 | Gurjit Singh |  | Dr. Nitin Dange | Cranial |  |  | 100 | 7.56 | 0.44 |  | Yes |  | https://drive.google.com/open?id=1zl2eyCRX87aIlMTYuje8FjWGwHzlbzUh | MRI Dataset wasn’t as per requirements. To less space in OT to operate the mobile system. | Accuracy was good | https://drive.google.com/open?id=15R4XocJuJhgSrXZO2p9sBtzi9dEuqQRe | Yes | NA | NA | There should be wheels attached to tripods stand to move the assembled system easily which could be locked for stable parking.
2024-09-28 21:52 | trushal@rxoom.com | Trushal Gulhane | BMHRC | Bhopal | NS TRL 01 MF 101 |  | V1.2 | Biopsy needle provided |  |  | https://drive.google.com/open?id=1LBzS3WSDe3UMSIVAGhH8hmFxIe3s9F4F | 2024-09-28 | Trushal Gulhane |  | Dr sandeep sorte | Biopsy, Any (Cranial, Full body) |  |  | 0 | 4.5 | 1.2 |  | Yes |  | https://drive.google.com/open?id=1gpGIFGhGhOVcLSbiYid0Em-HbPh-9AWL | NA | Yes | https://drive.google.com/open?id=1VRLpHS9cKniZCuNl914LbGw2FAmu8uIZ | Yes | NA | NA | NA
2024-10-01 | rohan@rxoom.com | Rohan Ghosh | Glenfield Mallareddy Hospital | Hyderabad | NS TRL 01 SE 101 |  | R 1.2 | NA |  |  | https://drive.google.com/open?id=1qPZT6Pr5jg909zI39-FAEDRyHLoc9TZX | 2024-09-27 | Rohan Ghosh |  | Dr Jalee; Kirmani & Dr T Satish Kr | Cranial |  |  | 0.05 | 5.56 | 0.74 |  | Yes |  | https://drive.google.com/open?id=1NuyRV_EjXAg7mDyoYF4L96PLV68elIXg | Initial technical issues with the data testing model weren't rotating , but after restarting the application once, the issue wasn't repeated. | 1.During the demo, Dr. Satish successfully performed fiducial registration, despite minor difficulties with surface registration due to patient hair. The system demonstrated decent accuracy. 2. Dr. Jaleel Kirmani utilized the virtual stylus option to confirm trajectory planning. 3. Post-surgery, Dr. Jaleel Kirmani reported precise accuracy and exact trajectory alignment, expressing satisfaction with our system's performance. 4. Dr. T Satish noted a significant reduction in surgical time using our system. | https://drive.google.com/open?id=1Iz8l1nDnZChr0XHWGRmW4FVaxx7ddwyd | Yes | NA | NA | NA
2024-10-05 11:30 | gurjit758.232@gmail.com | Suraj Garad | Surya Syhadri hospital | Pune | NS TRL 01 MS 101 |  | R 1.2 | NA | https://drive.google.com/open?id=1YviDHVC6wabCD4sfxAewqSl8HBpPTSX6 | 10/5/0024 | Gurjit Singh | Dr. Ranjit Deshmukh | Cranial |  |  | 100 | 2.59 | 0.63 | https://drive.google.com/open?id=1jdVEJ3bCuCCrQMKbKDJi8apqUQWZMtSw | NA | Accuracy was good even with any fixcation | https://drive.google.com/open?id=1vcQ8ReeFX9O0VuYVFQC8bsIkQyKc1jLf | Yes |  |  | NA | NA | NA | NA
2024-10-03 21:10 | hemarshitha1802@gmail.com | Hemarshitha H. S. | HCG hospital | Bengaluru | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1TaR6XQIm6MqGQIRCXx2-H1nl8-derRS- | 2024-10-03 | Hemarshitha H S | Dr Amresh | Cranial |  |  | 0.42 | 5.08 | 0.93 | 0.42 | 5.08 | 0.93 |  | Yes |  | https://drive.google.com/open?id=1yAxDQvOqDyyhfFWYVuYcxGZqRTMI8bzw | NA | Surgeon informed that the accuracy remained perfect throughout the surgery, mentioning that it was extremely helpful in this case. | https://drive.google.com/open?id=1MEx9qUWzxJxwX1LDgnf2N2fDp_QYNv4p | Yes | NA | NA | NA
2024-10-05 11:30 | gurjit758.232@gmail.com | Suraj Garad | Surya Syhadri hospital | Pune | NS TRL 01 MS 101 |  | R 1.2 | NA | https://drive.google.com/open?id=1YviDHVC6wabCD4sfxAewqSl8HBpPTSX6 | 10/5/0024 | Gurjit Singh | Dr. Ranjit Deshmukh | Cranial |  |  | 100 | 2.59 | 0.63 | 100 | 2.59 | 0.63 |  | Yes |  | https://drive.google.com/open?id=1jdVEJ3bCuCCrQMKbKDJi8apqUQWZMtSw | NA | Accuracy was good even with any fixcation | https://drive.google.com/open?id=1vcQ8ReeFX9O0VuYVFQC8bsIkQyKc1jLf | Yes | NA | NA | NA
2024-10-08 11:09 | rohan@rxoom.com | Rohan Ghosh | Rainbow Hospital | Hyderabad | NS TRL 01 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1kJvrid_shGcH5J8bxSFJ8xLPWbDRJ2BE | 2024-10-05 | Rohan Ghosh | Dr Vishakha Karpe | Cranial |  |  | 0.09 | 2.74 | 0.78 | https://drive.google.com/open?id=182x2JyZN7hZjR9lCggkRTTo7QKBGzkdL | Dr. Vishakha and Dr. Sarang Bakle collaborated on a medical case. Dr. Vishakha initiated the registration process, demonstrating proficiency in the technique. However, a minor setback occurred when the brother inadvertently moved a crucial reference point after the patient was draped. This disruption led to a loss of accuracy in the registration. | The NeuroSaphireX system demonstrated strong performance during the surgical procedure at Rainbow Hospital. The team found the system to be user-friendly, with intuitive navigation and accurate guidance. The registration process was straightforward, and the instruments were ergonomic and comfortable to use. The system's accuracy in performing the right pterional craniotomy was particularly impressive, achieving a margin of error of 0.7mm. Overall, the feedback from the surgical team indicates that the NeuroSaphireX system has the potential to significantly enhance surgical outcomes and efficiency .The team particularly appreciated the skull tr featackerure, which they found to be highly effective. However, they did identify a need for improved navigation features for endoscope procedures. The team also suggested that integrating enhanced visualization capabilities would further enhance the system's effectiveness. Overall, the feedback indicates that the NeuroSaphireX system has the potential to significantly improve surgical outcomes, but there is room for further development and refinement, particularly in the areas of endoscope navigation and visualization. | https://drive.google.com/open?id=1L02_bLiCRwT5TQS7HnpyGMWE7BlqunIv | Yes |  |  | NA | NA | NA | Build a display cover | Doctor asked for endoscopic spine navigation
2024-10-09 19:37 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Brains hospital | Bengaluru | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=11_EhB8x8e5gTnvvp4BccFdz85xF3aY6F | 2024-10-09 | Hemarshitha H S | Dr Raj Kumar & Dr Venkataramana | Cranial |  |  | 0.22 | 2.98 | 0.28 | https://drive.google.com/open?id=11vpyjkdNYQTHOrvKhmba2ZFC6Rf0ughv | NA | It would be great if we can fix the arm to sugita head frame. | Kaagaz_20241009_193143591392 (1) - Hemarshitha Hemarshitha.pdf | Yes |  |  | NA | NA | NA | NA | Fixation of arm directly to sugita head frame.
2024-10-10 22:39 | hemarshitha1802@gmail.com | Hemarshitha H. S. | HCG Hospital | Bengaluru | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1f0WRehwlvoYItA1EKt9Su07bztQ4Tr9F | 2024-10-10 | Hemarshitha H S | Dr Amresh | Cranial |  |  | 0.17 | 5.17 | 0.98 | https://drive.google.com/open?id=1itszqp_X--M4tEchkjF4DEI6UcQppvrd | Initially the camera was not detecting the reference, but after restarting the system, it worked. | NA |  | Yes |  |  | NA | NA | NA | NA | NA
2024-10-16 09:35 | garadsuraj567@gmail.com | Suraj Garad | DR L H HIRANANDANI HOSPITAL | Mumbai | NS TRL 01 MS 101 |  | R1. 2 | Long probe | https://drive.google.com/open?id=14ag2UpHjahoNWOBMBODPoMcisKjzQ9QB | 2024-10-15 | Suraj Garad & Janhvi Barhate | Dr. Sachin Giri | Cranial |  |  | 0.23 | 5.83 | 0.71 | https://drive.google.com/open?id=1BM1wvhTvJGy9GxIHpdGT72uRjnT-mF6J | Because of mobile system tripod we have to lift the camera and adjust it because in some angle probe is not getting defected. | NA |  | Yes |  |  |  | NA | NA | NA | NA
2024-10-20 23:17 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Bhagwan Mahaveer Jain hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1tuqQuzelOG_UVR8ZHHVnAt7KORqYTXAN | 2024-10-19 | Hemarshitha H S | Dr Krishnaprasad, Dr Sharan Srinivas & Dr Sasidhar | Cranial |  |  | 0.17 | 4.81 | 0.48 | https://drive.google.com/open?id=1s4n-5DaBkpw_UA4sHh20MXofAWp8RJ7X | NA | 1. Dr. Sasidhar noted that the registration process was simpler compared to Medtronic, with a fast workflow and no lag. 2.Dr. Krishnaprasad was pleased with the system's accuracy and registration process. 3.Dr. Sharan Srinivas appreciated the use of Apple products, found the system performance very helpful during the surgery. | https://drive.google.com/open?id=1D_AWpY03vnBt-yhyZBRc0FriiSazhIbM | Yes |  |  | NA | NA | NA | NA | NA
2024-10-23 21:31 | rohan@rxoom.com | Rohan Ghosh | Rainbow Hospital | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1mgKK0od1uflHrsK73hgI-ja4z6yORwIP | 2024-10-15 | Rohan Ghosh | Dr Praneeth Kokkula | Cranial |  |  | 0.2 | 8.34 | 1.5 | https://drive.google.com/open?id=1k00EoUtXs13Gz1NDhxDjEtSa3U5UYo0m | At starting when we were doing registration, we face one issue. The problem was, that, after completing the registration process our probe was showing at the frontal part of face. But when we redo the same process, system didn't show it again. | Software & Application upgrade required | https://drive.google.com/open?id=1N4_a0Ap8sSUz2lSPAYrAoOoWkYIIhER3 | Yes |  |  | At starting when we were doing registration, we face one issue. The problem was, that, after completing the registration process our probe was showing at the frontal part of face. But when we redo the same process, system didn't show it again. | NA | NA | As per doctor our system need lot of software & application update, as an example biopsy needle projection. | As per doctor our system need lot of software & application update, as an example biopsy needle projection.
2024-10-23 21:49 | rohan@rxoom.com | Rohan Ghosh | ESIC Hospital | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Biopsy Needle | https://drive.google.com/open?id=1PV7IsQnAYz0iHR1pwXp-mAmjxy9x6lXN | 2024-10-23 | Rohan Ghosh | Dr P Rajesh | Biopsy, Any (Cranial, Full body) |  |  | 0.08 | 6.75 | 0.93 | https://drive.google.com/open?id=1wQY0NTVEIVu3anruVtjOwqxDsMyWTmrA | During the surgery, the OT staff inadvertently disconnected the main power supply twice. | Doctor was asking for a extra display. But Dr. Rajesh, the Unit HOD, expressed high satisfaction with the system's performance and noted that the software was easy to learn and operate. He showed interest in using the system for future surgeries, including spine navigation cases. | https://drive.google.com/open?id=1nRyNCUwMzSUe1-j_GeFdTr0R6pjRNtnW | Yes |  |  | NA | NA | NA | After accidently system has been shut down, when I restart the system , then I notice that the color of trac plan has been changed. | Doctor was asking for another display should be integrated with our system & he also demand that we should carry 4 universal trackers during demo.
2024-10-29 19:09 | hemarshitha1802@gmail.com | Hemarshitha H. S. | East Point Hospital | Bangalore | NS TRL 02 MF 204 |  | V1.02 | NA | https://drive.google.com/open?id=1rBsV9U85CAFgEwfHS4sV6Mrl_n3ksJm_ | 2024-10-28 | Hemarshitha HS | Dr. Mahendra | X-Ray Spine |  | 15 | 0.04 | 00 | 0.57 | https://drive.google.com/open?id=1idOZVPrVH3EePvA3yLxg8yBoukR-Ymef | NO | Feedback from the surgeon: Initially, the accuracy was perfect from the L3-L5 region, but later in the L5-S1 region, there was a slight accuracy error. However, the doctor understood the situation and expressed happy with the system's performance. |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-11-04 10:50 | Janhavi@rxoom.com | Janhvi Barhate | Magnum Heart Institute Hospital | Nashik | RX - MT 0923 - 01 |  | V.1.02 | NA | https://drive.google.com/open?id=154hIj6tzuSGyc6xrKSxDIuBoiRdsdTZj | 2024-10-29 | Janhavi Barhate | Dr. Rahul Birla | Cranial |  |  | 0.75 | 0.75 | 0.75 | https://drive.google.com/open?id=1i6mo8nz4nMeQU581S0hgGxuMu8PBogjR | NA | Registration process is simple. Located tumor precisely, happy with the biopsy. |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-11-04 10:59 | trushalgulhane1547@gmail.com | Suraj Garad | Neuron hospital | Nagpur | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=14tUKiu_EtoKurII1pQB5IyXUyVFhtnnl | 2024-10-30 | Trushal Gulhane | Dr. Pramod giri and Dr. Sanjok gajbhiye | Cranial |  |  | 0.2 | 2.09 | 0.7 | https://drive.google.com/open?id=1UyeUejKnMw4hqje7l8AAfWmx9ZAiLt4X | NA | Good feedback |  | Yes |  |  | NA | NA | NA | Ask to train their OT staff | NA | [13 empty] | Noted
2024-11-04 19:31 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Sri Sathya Hospital Bangalore | Bangalore | NS TRL 02 MF 204 |  | V.02 | NA | https://drive.google.com/open?id=13yCTJRx41QWxWlJ8Jog_MBWHZ3c_XXCo | 2024-10-30 | Hemarshitha HS | Dr. Saritha Dr. Anand | Cranial |  |  | 0.19 | 2.98 | 0.78 | https://drive.google.com/open?id=1U-9QLUALdecrlay1qZA6fGipOuPcn6O7 | From the initial stage only accuracy issues happened. Even tried four times registration but accuracy was not good and even though Surface accuracy less than 1 mm, and fiducial accuracy less than 5 mm but, after registration process the probe was showing a major error like it's already inserted. Even though I tried to change the galions and probes still the problem was same. | Surgeon was not satisfied and she was angry. |  | No | NA | No Idea because it's happening in that hospital only. | NO | NO | NO | NA | NA | [13 empty] | Noted
2024-11-07 00:39 | rohan@rxoom.com | Rohan Ghosh | ESIC Hospital | Hyderabad | NS TRL 01 MS 101 |  | 1.02 | Universal tracker | https://drive.google.com/open?id=1uPjwi9GyvYEm5wrP41pyvxeK1KUU7gu4 | 2024-11-29 | Rohan Ghosh | Dr. P Rajesh | CT-Spine |  | 1000 | 0.15 | 2.31 | 2.84 | https://drive.google.com/open?id=1tLiQloQ1soOAAgJlqXta5WTxTzME6A-Q | 1. During the procedure, the CT spine clamp and reference markers were inadvertently moved, necessitating a second registration attempt. However, the subsequent registration did not achieve the desired level of fiducial accuracy. 2.The surgeon expressed interest in using a navigation system and calibrating personal instruments, such as pedicle finders. Unfortunately, compatibility limitations prevented the integration of these devices with the existing system. | Dr. P. Rajesh performed a CT spine and stereotaxic craniotomy biopsy surgery at ESIC Hospital, Hyderabad, using the NeuroSaphireX Navigation System. The procedure achieved an accuracy of 2.3 mm. Dr. Rajesh provided the following feedback: 1. Positive Aspects: Dr. Rajesh appreciated the ease of loading patient data, simple registration process, and user-friendly interface, marking these aspects as “Agree” or “Strongly Agree.” 2. Concerns: Accuracy was noted as a concern, with additional feedback requesting improvements, especially for spine cases. 3.Feature Suggestions: Dr. Rajesh recommended adding a pedicle finder with a tracker to the system to enhance utility. 4. System Limitation: The tracker component was identified as less useful. 5. Overall Satisfaction: Dr. Rajesh expressed general satisfaction, noting the system's ease of learning and potential for additional surgeries, including spine navigation. This feedback will be instrumental in guiding improvements to optimize system performance and meet specific clinical requirements. | https://drive.google.com/open?id=1l5QpFQqx2JVPASCVJF426uaYnH4nV6EF | Yes |  |  | 1. The surgeon wants the ability to change slices while in navigation mode when the probe is visible on the screen. 2.The surgeon requests on-the-fly calibration for surgical tools. | NA | NA | Develop Pedicle finder, which should be pre calibrated & should be navigated too. | Dr. P. Rajesh performed a CT spine and stereotaxic craniotomy biopsy surgery at ESIC Hospital, Hyderabad, using the NeuroSaphireX Navigation System. The procedure achieved an accuracy of 2.3 mm. Dr. Rajesh provided the following feedback: 1.Positive Aspects: Dr. Rajesh appreciated the ease of loading patient data, simple registration process, and user-friendly interface, marking these aspects as “Agree” or “Strongly Agree.” 2. Concerns: Accuracy was noted as a concern, with additional feedback requesting improvements, especially for spine cases. 3. Feature Suggestions: Dr. Rajesh recommended adding a pedicle finder with a tracker to the system to enhance utility. 4. System Limitation: The tracker component was identified as less useful. 5. Overall Satisfaction: Dr. Rajesh expressed general satisfaction, noting the system's ease of learning and potential for additional surgeries, including spine navigation. This feedback will be instrumental in guiding improvements to optimize system performance and meet specific clinical requirements. | [13 empty] | Noted
2024-11-07 12:02 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | R 1.2 | NA | https://drive.google.com/open?id=1sIv7se334Lpj9AoaUT1F_RhYSROtz75O | 2024-11-06 | Gurjit Singh | Dr. Pawan Goal | Cranial |  |  | 0.09 | 4.36 | 0.85 | https://drive.google.com/open?id=1sLKNCnQUytaXrGR-QrNqwS4FmjIL5o8- | Model takes alot of time to generate. And T1 dataset which is our primary dataset wasn’t generating good and smooth model and taking a lot of try and error. While T2 generates good and smooth model in less tries. | Accuracy is good |  | Yes |  |  | NA | NA | Porter was not available at RML hospital and dealer person used innova Cab and dissembled the system base to easy system transport. While unloading the base from cab at hospital fall on his foot thumb as base was heavy. | System desk or flap border should be more round or curvy and smooth look, it should not be in square shape as it give look of square metal box look. | Option of DBS, dual monitor, real-time model generation. | [13 empty] | Noted
2024-11-11 18:07 | gurjit758.232@gmail.com | Gurjit Singh | Paras Hospital | Gurgaon | NS TRL 01 MF 101 |  | R1.2 | NA | https://drive.google.com/open?id=1ucBUfMOB7Q1TZsNJl9_8GzFtA9y51zXV | 2024-11-11 | Gurjit Singh | Dr Manmohan Singh and Dr Ravi | Cranial |  |  | 100 | 2.73 | 0.62 | https://drive.google.com/open?id=1SSDL6fwJ3Pl2-Wgtu_soFwG5Tt08GYFy | NA | Good accuracy |  | Yes |  |  | NA | NA | Took a lot of force to move the trolley and moving one side while dragging or moving trolley | NA | Need Cross section in navigate mode to identify and mark midline easily | [13 empty] | Noted
2024-11-11 22:12 | trushalgulhane1547@gmail.com | Trushal Gulhane | Neuron hospital | Nagpur | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1M1pMIaJ6fRt2Q43jTtOew0X9IxLHJ9YN | 2024-11-11 | Trushal Gulhane | Dr. Pramod giri and Dr. Sanjok gajbhiye | Cranial |  |  | 0.2 | 4.3 | 1.04 | https://drive.google.com/open?id=1hWdtx6TXwsRQIX_hvLl53Qtb633x7EgC | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-11-12 23:41 | hemarshitha1802@gmail.com | Hemarshitha H. S. | East point hospital | Bengaluru | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1HhkYa4OgVwfKusDGZs-XqsdK20F5T_lA | 2024-11-12 | Hemarshitha H S | Dr Mahendra | Cranial |  |  | 0.13 | 4.51 | 0.33 | https://drive.google.com/open?id=1HLVKuCcwugiIHSPS7__P5tw-GrYpFdhv | NA | Surgeon was highly impressed with the system's performance, calling it a commendable Indian innovation. |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-11-16 20:32 | gurjit758.232@gmail.com | Gurjit Singh | Paras Hospital | Gurgaon | NS TRL 01 MF 101 |  | R1.2 | NA | https://drive.google.com/open?id=13N1ShSSBMcPU72yfcmC0tAOIsEIDxSpS | 2024-11-13 | Gurjit Singh | Dr Manmohan Singh and Dr Ravi | Cranial |  |  | 0.14 | 2.48 | 1.18 | https://drive.google.com/open?id=1YNvY19eFbHixN1pIBSsaU9oa8tgabtpR | Accuracy loosed before dura opening. Which could be due to arm movement due to excessive weight on arm. Or retractor arm force on arm. | Arm should not move even using someone use whole body weight on the arm. |  | No | Accuracy should be maintained through out the surgery but accuracy loosed a before opening the dura | I don’t have any certain opinion. But one factor could be moving of arm with excessive weight on arm. | NA | NA | While moving trolley, it get moving any direction not the direction we want. It took lot of force to handle the trolly | Light weight trolly | NA | [13 empty] | Noted
2024-11-16 21:40 | gurjit758.232@gmail.com | Gurjit Singh | Paras Hospital | Gurgaon | NS TRL 01 MF 101 |  | R1.2 | NA | https://drive.google.com/open?id=1VpOYzwlIS28Ti_2GyWrS9ieQamVk_1-1 | 2024-11-14 | Gurjit Singh | Dr Raghav | Cranial |  |  | 0.1 | 3.33 | 0.56 | https://drive.google.com/open?id=158emyC1jcQbKuPcsuCukY3gHsYsPCGpr | NA | Good Accuracy |  | Yes |  |  | NA | Biopsy needle Stopper by CSSD or sister who handle sterile instruments in operation | Trol | NA | NA | [13 empty] | Noted
2024-11-25 11:33 | janhavijb99@gmail.com | Faiz Khan | Deenanath Mangeshkar Hospital | Pune | Other | DMH's System | R1.2 | universal table adaptor, Extension arm, Biopsy tool, Biopsy traker, Screw driver, 2 biopsy arms. | https://drive.google.com/open?id=1w65egJwbbxRKtnhtetxkT2GQk2XjAeWK | 2024-11-23 | Janhavi Barhate | Dr. Akshay Gupte | Biopsy, Any (Cranial, Full body) |  |  | 0.82 | 2.49 | 0.82 | https://drive.google.com/open?id=16TXigO5BilVwWVJ4gahgIt7gIS1FlE-q | NA | Accuracy is good. Need to improve biopsy arm. Need to improve Needle tracker's working. |  | Yes |  |  | NA | Na | NA | Need to improve the movability of biopsy arm | NA | [13 empty] | Noted
2024-11-27 18:10 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS Delhi | Delhi | NS TRL 01 MF 101 |  | R1.2 | NA | https://drive.google.com/open?id=1_OqwIRyULNMVwfhmoGtbTKe7nH-Q7R7x | 11/19/0024 | Gurjit Singh | Dr Amol | Cranial |  |  | 0.07 | 4.48 | 0.57 | https://drive.google.com/open?id=15y_ZLWe4kDOKQK6HUSMoP39EkgjVfVth | Not able to show brain tissues clearer like HRS system | Good accuracy add Brain window feature |  | Yes |  |  | NA | NA | NA | NA | Work on Brain window function. | [13 empty] | Noted
2024-11-27 18:21 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS Delhi | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1fWmcR9jKilnjgb4rfZqmCy6HiHC4cG5E | 11/20/0024 | Gurjit Singh | Dr. Amol | Cranial |  |  | 0.11 | 3.74 | 0.58 | https://drive.google.com/open?id=1bI_k-9V4J9FkENSfJjWiJsQa2g5HcJYP | NA | NA |  | Yes |  |  | NA | Na | Na | NA | NA | [13 empty] | Noted
2024-11-27 18:31 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS Delhi | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=18Pqnvhupu5dH2L9B9Ta990k6HDN_hdDI | 11/20/0024 | Gurjit Singh | Dr Amol | Cranial |  |  | 0.11 | 2.54 | 0.49 | https://drive.google.com/open?id=1SRYAKJZPFaQ0RBPjXm2BLxjPsabjbbky | NA | Good accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-11-27 18:40 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS Delhi | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1Ip8OzbxX7NsWsyEnGw0c0in9-f3Vm2Hr | 11/22/0024 | Gurjit Singh | Dr Vishal | Cranial |  |  | 100 | 4.26 | 0.52 | https://drive.google.com/open?id=1-4MaDgRvOf8IoR169GqFeVWsMycTj89f | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-14 22:12 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1O_aH10onigApPVO_0krvbfNdP9xiPVzS | 11/25/0024 | Gurjit Singh | Dr Amol | Cranial |  |  | 100 | 2.45 | 0.62 | https://drive.google.com/open?id=1Q9paalr3Go4Ixw_t9zoqaO1oRmnB1ewN | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1ob-tJtioDLRDL6rvTBxoWlDOUbqsTcO6 | [12 empty] | Noted
2024-11-27 18:53 | gurjit758.232@gmail.com | Gurjit Singh | AIIMS Delhi | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1Utxc6urCvGj_2ubmQcI-o9Php990ixIt | 11/25/0024 | Gurjit Singh | Dr Amol | Cranial |  |  | 100 | 3.72 | 0.73 | https://drive.google.com/open?id=1CU50Pgy6Buf3_IyvkhB11Z0Bvf5PnK-h | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-02 21:50 | rohan@rxoom.com | Rohan Ghosh | NIMS | Rohan Ghosh | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1P7dvmHevUXKJScObqGGM_G_5Wc7JKL4B | 2024-11-18 | Rohan Ghosh | Dr Tirumala Reddy | Cranial |  |  | 0.27 | 8.52 | 0.93 | https://drive.google.com/open?id=1JK3d-ByhxN7BTL7DjtmSdB6IIzvUT9vq | During the surgery, a significant cerebrospinal fluid (CSF) leak caused brain shifting. As a result, Dr. Tirumala Reddy remarked that navigation accuracy becomes compromised when brain shifting occurs. However, he acknowledged the system's performance prior to the brain shift and expressed satisfaction with its initial accuracy. | 1.System Performance: The NeuroSaphireX system demonstrated high reliability and ease of use during the procedure. 2.Data Management: Patient data was loaded without any complications. 3.Registration Process: The registration process was completed efficiently and maintained simplicity throughout. 4.Navigation Accuracy: The navigation system provided satisfactory accuracy, ensuring confidence in its functionality. 5.Ergonomics and Usability: The instruments were ergonomic, and the User Interface (UI) was intuitive and user-friendly. | https://drive.google.com/open?id=1DgMusPMK0jQS2g16EkIvRyWEcOtaLWWz | No | During the surgery, a significant cerebrospinal fluid (CSF) leak caused brain shifting. As a result, Dr. Tirumala Reddy remarked that navigation accuracy becomes compromised when brain shifting occurs. However, he acknowledged the system's performance prior to the brain shift and expressed satisfaction with its initial accuracy. | During the surgery, a significant cerebrospinal fluid (CSF) leak caused brain shifting. As a result, Dr. Tirumala Reddy remarked that navigation accuracy becomes compromised when brain shifting occurs. However, he acknowledged the system's performance prior to the brain shift and expressed satisfaction with its initial accuracy. | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-02 21:52 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Biopsy Needle | https://drive.google.com/open?id=1-0Ocupo8g6ggWto1IX35Bt-bSuz7l2cO | 2024-11-20 | Rohan Ghosh | Dr. Ramamath | Biopsy, Any (Cranial, Full body) |  |  | 0.13 | 2.94 | 1.1 | https://drive.google.com/open?id=1Fh2wMQkuh1Yi0yD5ACTS3IpssH08NPZQ | 1. The surgical team encountered issues with the track plan sequence, noting discrepancies between different planes. This inconsistency caused dissatisfaction with the software's current planning module. 2. The navigation camera intermittently failed to recognize the reflecting balls, particularly those attached to the biopsy needle, leading to delays. 3. During the procedure, the biopsy needle was not displayed in real-time, causing inaccuracies in the navigation. The surgeon remarked that the displayed data did not align with the actual position of the needle. | Not given |  | No | 1. The surgical team encountered issues with the track plan sequence, noting discrepancies between different planes. This inconsistency caused dissatisfaction with the software's current planning module. 2. The navigation camera intermittently failed to recognize the reflecting balls, particularly those attached to the biopsy needle, leading to delays. 3. During the procedure, the biopsy needle was not displayed in real-time, causing inaccuracies in the navigation. The surgeon remarked that the displayed data did not align with the actual position of the needle. | 1. Dr. Ramanath emphasized the need for a more robust and complete biopsy project software sequence. He requested the development and installation of an updated module to address these issues. | NA | Biopsy needle | NA | According to Dr. Ramanath emphasized the need for a more robust and complete biopsy project software sequence. He requested the development and installation of an updated module to address these issues. | Dr. Ramanath emphasized the need for a more robust and complete biopsy project software sequence. He requested the development and installation of an updated module to address these issues. | [13 empty] | Noted
2024-12-02 21:48 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1qd-wP2ACkUFh-48wq5e_rAezyuerLwl9 | 2024-11-11 | Rohan Ghosh | Dr Vamsi | Cranial |  |  | 0.1 | 3.65 | 0.48 | https://drive.google.com/open?id=1DMJ8Tsg89HpSpohK8vIWliPL7Cnw3HlB | NA | The surgeon described the navigation system's accuracy as "beautiful." | https://drive.google.com/open?id=1eVbkfRwpQbDifr3EtYn6O736SOjyjM0i | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-03 00:48 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Skull Tracker | https://drive.google.com/open?id=1nJC7S8GVE0BM3ET9SRQff_tM6eH6_oHv | 2024-11-13 | Rohan Ghosh | Dr. Vamsi | Cranial |  |  | 0.1 | 4.55 | 0.88 | https://drive.google.com/open?id=1d0yhvD-t7zjz1adWQFDsLVrHRveEUyS3 | NA | At the conclusion of the surgery, both Dr. Vamsi Yerramneni and Dr. Murli Krishna expressed their high level of satisfaction with the system. They acknowledged it as a superior navigation system, particularly in terms of accuracy and reliability. | https://drive.google.com/open?id=1SQqp9juZ3nFskmFW_qzNabq7Prmcxri1 | Yes |  |  | NA | NA | NA | NA | 1. Dr. Murli Krishna, an integral part of the surgical team, exhibited high confidence in the navigation system from the beginning of the procedure. 2.At the conclusion of the surgery, both Dr. Vamsi Yerramneni and Dr. Murli Krishna expressed their high level of satisfaction with the system. They acknowledged it as a superior navigation system, particularly in terms of accuracy and reliability. | [13 empty] | Noted
2024-12-03 00:54 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Skull Tracker | https://drive.google.com/open?id=1URzE5p9Y7AELtfmljHUL-IIyIUSclsim | 2024-11-14 | Rohan Ghosh | Dr. Vamsi | Cranial |  |  | 0.15 | 3.5 | 0.45 | https://drive.google.com/open?id=1xhsCLKbBgFN3VbWWFvmLrs5pn3walddt | NA | The surgery was successfully completed, with our navigation system demonstrating exceptional performance. The positive feedback from the surgical team highlights the system's potential for broader applications in complex procedures. | https://drive.google.com/open?id=1ejSCUo_f0jN4OptrhCU3oRni_lVEsK7L | Yes |  |  | NA | NA | NA | NA | At the conclusion of the surgery, both Dr. Vamsi Yerramneni and Dr. Murli Krishna expressed their high level of satisfaction with the system. They acknowledged it as a superior navigation system, particularly in terms of accuracy and reliability. | [13 empty] | Noted
2024-12-03 01:06 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Biopsy Needle, Skull tracker | https://drive.google.com/open?id=17G9o8LinM6Mne0aM_XsKjxzTb1POmTF7 | 2024-11-18 | Rohan Ghosh | Dr. Vamsi | Biopsy, Any (Cranial, Full body) |  |  | 0.16 | 3.17 | 0.76 | https://drive.google.com/open?id=1knSzypkat7eMklkncZR77ms-dNGnkLPj | The surgeon marked the track plan on various planes using the system. However, Dr. Vamsi expressed concerns regarding the sequence of operations within the software, indicating that it did not meet his expectations in this specific case. | The navigation system demonstrated solid performance during the brain biopsy, particularly with registration accuracy and procedural guidance. Feedback regarding the software's sequence will be evaluated for potential improvements to enhance user satisfaction. The successful use of the skull tracker and biopsy needle combo reflects the system’s capability for such procedures. | https://drive.google.com/open?id=1Lc47eWj4Mzm3lA61fq9boMP39tFc8d41 | Yes |  |  | The surgeon marked the track plan on various planes using the system. However, Dr. Vamsi expressed concerns regarding the sequence of operations within the software, indicating that it did not meet his expectations in this specific case. | NA | NA | Doctor has requested the development of a biopsy projection software to be integrated into the planning software sequence. This software will enable doctors to accurately plan and confirm the trajectory or track plan for biopsies, enhancing patient safety. | Doctor has requested the development of a biopsy projection software to be integrated into the planning software sequence. This software will enable doctors to accurately plan and confirm the trajectory or track plan for biopsies, enhancing patient safety. | [13 empty] | Noted
2024-12-03 01:14 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | Biopsy Needle, Skull Tracker | https://drive.google.com/open?id=11DpvYEZEbUVuvf7KVhr0XWSB1hF8YOoZ | 2024-11-20 | Rohan Ghosh | Dr. Murli Krishna | Biopsy, Any (Cranial, Full body) |  |  | 0.07 | 5.41 | 0.73 | https://drive.google.com/open?id=1vhQAW0jl5p61O3IFh4h_Fi5dP95RwgsT | The tumor lesion was considerably large, enabling the surgeon to collect four biopsy samples with ease. However, during navigation, the track plan occasionally did not display. This required reverting to setup mode and reselecting the track plan multiple times to enable visibility in navigation mode. | Despite the minor inconvenience in track visibility, Dr. Murli Krishna and his team expressed overall satisfaction with the performance of the navigation system, commending its accuracy and reliability during the procedure. | https://drive.google.com/open?id=11okYOmKmYfFErTGDiQoZPfWVXvRfT_oO | Yes |  |  | NA | NA | NA | Feedback regarding the intermittent track visibility issue will be addressed to enhance the system's usability. | Feedback regarding the intermittent track visibility issue will be addressed to enhance the system's usability. | [13 empty] | Noted
2024-12-03 01:20 | rohan@rxoom.com | Rohan Ghosh | NIMS | Rohan Ghosh | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1MRf321fB3Hd8gXO6Wkmgy1BLRdDMIsYS | 2024-11-21 | Rohan Ghosh | Dr. Vamsi | Cranial |  |  | 0.27 | 2.85 | 0.97 | https://drive.google.com/open?id=1x1I-1pt5-3wTOT0ktBkzQUuI08l-1aFk | NA | The navigation system maintained consistent accuracy throughout the entire procedure, ensuring a smooth and precise surgical workflow. | https://drive.google.com/open?id=1h0y9skCPU3ZcrkPRo8AJrI32puicBklE | Yes |  |  | NA | NA | NA | NA | The surgery was successfully completed, with our navigation system demonstrating exceptional performance. The positive feedback from the surgical team highlights the system's potential for broader applications in complex procedures. | [13 empty] | Noted
2024-12-03 01:34 | rohan@rxoom.com | Rohan Ghosh | NIMS | Hyderabad | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=102H5eK0sGqPSmhbHP1GYMT4jiKbJ7gbE | 2024-11-23 | Rohan Ghosh | Dr. Vasundhara | Cranial |  |  | 0.12 | 2.85 | 1.02 | https://drive.google.com/open?id=15dhkop_4jq2j6NKJTeUU-S6pw7wJANka | Post patient draping and pre-surgery, our navigation system displayed optimal accuracy. However, a critical incident occurred towards the end of the procedure. A junior surgeon, assisting the senior surgeon, inadvertently moved the navigation arm, compromising the system's accuracy at a crucial juncture. | NA | https://drive.google.com/open?id=1g3ELhBPaZan0PIPIsfQ9ip4W83-jroed | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-03 02:01 | rohan@rxoom.com | Rohan Ghosh | NIMS | Rohan Ghosh | NS TRL 03 SE 101 |  | 1.02 | NA | https://drive.google.com/open?id=1ghUpy-3gd9LOsnPAZ_s023uab_wlGAq3 | 2024-11-25 | Rohan Ghosh | Dr. Vamsi | Cranial |  |  | 0.18 | 7.58 | 1.65 | https://drive.google.com/open?id=1baGOrTRTqoBAbzAblCga_hQ70LDydzuS | During the surgical procedure, we encountered difficulties in acquiring accurate surface navigation points due to the patient's thick, curly hair. Additionally, as the surgeon was unfamiliar with our navigation system, the process of acquiring surface points was not as smooth as anticipated. Furthermore, a temporary slowdown in the navigation system software hindered the efficient placement of the four fiducial points, making this task more challenging than usual. | 1. Registration Accuracy It was observed that registration accuracy in certain patient positions requires improvement. This limitation has been noted as a potential challenge during specific procedures and may impact the overall precision of the system. 2. Hardware and Software for Biopsy Procedures The current hardware and software used for biopsy procedures were found to require enhancements. These improvements are essential to increase the accuracy and streamline the usability of the system, thereby optimizing performance during critical interventions. 3. Enhancement for Specific Registrations Particular attention needs to be given to the registration processes for scenarios where the patient's head is not fixed. Adjustments in the system are recommended to ensure seamless performance and reliability in such cases. | https://drive.google.com/open?id=1gzaDcX1sre_vLWZkGRBxffgnc9RktX5E | No | 1. Registration Accuracy It was observed that registration accuracy in certain patient positions requires improvement. This limitation has been noted as a potential challenge during specific procedures and may impact the overall precision of the system. 2. Hardware and Software for Biopsy Procedures The current hardware and software used for biopsy procedures were found to require enhancements. These improvements are essential to increase the accuracy and streamline the usability of the system, thereby optimizing performance during critical interventions. 3. Enhancement for Specific Registrations Particular attention needs to be given to the registration processes for scenarios where the patient's head is not fixed. Adjustments in the system are recommended to ensure seamless performance and reliability in such cases. | During the surgical procedure, we encountered difficulties in acquiring accurate surface navigation points due to the patient's thick, curly hair. Additionally, as the surgeon was unfamiliar with our navigation system, the process of acquiring surface points was not as smooth as anticipated. Furthermore, a temporary slowdown in the navigation system software hindered the efficient placement of the four fiducial points, making this task more challenging than usual. | NA | NA | NA | A. Improve Registration Accuracy: 1. Invest in research and development to enhance the registration algorithms. 2. Explore alternative registration methods that are less sensitive to hair and scalp variations. 3. Provide clear guidelines and training to surgeons on best practices for patient preparation to optimize registration accuracy. B. Reduce System Reset Frequency: 1. Develop algorithms that can adapt to variations in patient preparation, minimizing the need for system resets. 2. Implement features that allow for partial re-registration to correct minor inaccuracies without requiring a full system reset. C. Upgrade Software and Hardware: 1. Regularly update the system's software to incorporate the latest advancements in navigation technology. 2. Consider upgrading the hardware components to improve processing speed and overall performance. | 1. Invest in research and development to refine the accuracy of registration in varied patient positions. 2. Upgrade hardware and software capabilities for biopsy procedures to align with advanced clinical demands. 3. Focus on engineering solutions to improve the registration process in non-fixed head scenarios for broader applicability. | [13 empty] | Noted
2024-12-05 17:22 | gurjit758.232@gmail.com | Gurjit Singh | Vankateshwar Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1iF0u_Ud5hBU_wqo8QxX0NZdkznNgRJUu | 12/3/0024 | Gurjit Singh | Dr PK Sachdeva | Cranial |  |  | 0.11 | 4.55 | 0.43 | https://drive.google.com/open?id=1khxORnO_O_zF1-J9RJbt-Xi-LvlHUhxT | NA | Fast registration and good accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-05 17:34 | gurjit758.232@gmail.com | Gurjit Singh | Vankteshwar Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1cnYxu2oPoVfyh4z1-GwyqTwWKLiSVWsP | 12/3/0024 | Gurjit Singh | Dr P K Sachdeva | CT-Spine |  | 30 | 100 | 100 | 100 | https://drive.google.com/open?id=1oHJOXE--u1WnFGXMzJa30QXxPXLYaxyY | NA | The procedure should be more effortless. |  | No | Dr. was supposed to capture fiducials correctly on spine which he decided to capture. But Dr wasn’t able to capture the fiducials correctly. | Dr. was supposed to capture fiducials correctly on spine which he decided to capture. But Dr wasn’t able to capture the fiducials correctly. | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-05 17:44 | gurjit758.232@gmail.com | Gurjit Singh | Vankteshwar Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1PD9APemihkZvgvwlxe-4UCpZCX_kA4T0 | 12/4/0024 | Gurjit Singh | Dr. P K Sachdeva | CT-Spine |  | 30 | 100 | 4.05 | 100 | https://drive.google.com/open?id=1C-t4-8pJPlM519GBSiidNLVS5fMYPH1- | Software crash after completion of surface registration. | Keep redoing make them impatient |  | No | After surface point collection navigation should keep working and Dr use use the navigation but after completing of surface point software crash. While I restarted the software Dr got impatient and removed the references. | After surface point collection navigation should keep working and Dr use use the navigation but after completing of surface point software crash. While I restarted the software Dr got impatient and removed the references. | NA | NA | NA | NA | NA | [13 empty] | Noted
2024-12-13 16:01 | gurjit758.232@gmail.com | Gurjit Singh | Venkateswar Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1buHdElXaM34qtOd9msx8QUsUo3JnXYNq | 12/6/0024 | Gurjit Singh | Dr. P K Sachdeva | CT-Spine |  | 30 | 100 | 4.05 | 100 | https://drive.google.com/open?id=1dg3aQQkwysVW-KnNqfNur100XrH3nlXL | NA | NA |  | No | Software crash after the procedure of patient registration which should not be crash | Crashing of software and impatience of Dr | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1u6xWGu0KbkBdK3CSl-U5oYiHrjVErVTl | [12 empty] | Noted
2024-12-13 16:07 | gurjit758.232@gmail.com | Gurjit Singh | Venkateswar Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1ZaFhFvACTlP7WCsghM7Wvt-NbMuEGW2x | 12/9/0024 | Gurjit Singh | Dr. Manish | Cranial |  |  | 0.17 | 1.93 | 0.58 | https://drive.google.com/open?id=1KCNxqU8bkiN2shZyCptMlUFYTR1-AQkn | NA | Good Accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1Irdp2Wk6thTrJqs4KXhZTt6V-jIvVCsa | [12 empty] | Noted
2024-12-13 16:57 | garadsuraj567@gmail.com | Customer System, Sold on 19 Dec 2023 | Dr. Umesh Phalke | Pune | NS TRL 02 MF 203 |  | V1.0.2 | NA | https://drive.google.com/open?id=1S5SYXMVqdCjZe8b5L4o8revX6hgHDBJJ | 2024-12-12 | Suraj Garad | Dr. Umesh phalke | Cranial |  |  | 0.2 | 2.64 | 0.73 | https://drive.google.com/open?id=1cfya6qKjDsIaTP4obRzoY_mhfV2cfMqS | NO | They ask me to can we skip the surface registration from registration method |  | Yes |  |  | NA | NO | NA | NA | NA | https://drive.google.com/open?id=1Hrxrn2VBCNRKHP5Z7eQIGSBU99IJ6Zc9 | [12 empty] | Noted
2024-12-16 16:16 | gurjit758.232@gmail.com | Gurjit Singh | RML Hospital | Delhi | NS TRL 01 MF 101 |  | 1.2 | NA | https://drive.google.com/open?id=1DhAOhsHdgHDEVf4WF-7e3vtbID4PTkQ6 | 12/16/0024 | Gurjit Singh | Dr. Ajay Choudhary | Cranial |  |  | 0.06 | 2.35 | 0.55 | https://drive.google.com/open?id=11Zw4oFZLexKvu6gj8xjD7cSIW10dKtDE | NA | Good accuracy and system help guide in localization of tumor |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1hE761ockksOOAJRHAKmKDZVGPyO2duzD, https://drive.google.com/open?id=1muXO1ZnT6ItN1xxy_WG1cgBfPwlUythD, https://drive.google.com/open?id=1sxXQouEZHwW8q0_YNjIQsM_IuBMcxIWa, https://drive.google.com/open?id=1H5NQbNUxX2dVbYuhylivbj9emk1Y1d9k, https://drive.google.com/open?id=1NrSu4doPU2qew3qr-aPtOIkEcf_V3kQ4 | [12 empty] | Noted
2024-12-16 23:43 | rohan@rxoom.com | Rohan Ghosh | INK | Kolkata | NS TRL 01 SE 101 | NS TRL 01 SE 101 | 1.02 | NA | https://drive.google.com/open?id=1RC6An4ZhEHiMFORgbBiuPoFzM1hU2Y_2 | 2024-12-16 | Rohan Ghosh | Dr. Amit Kr Ghosh | Cranial |  |  | 0.08 | 4.48 | 0.64 | https://drive.google.com/open?id=10_FGIH9AQ-Z0R1_ry1q1E94ROU2EwWfq | NA | At the conclusion of the surgery, Dr. Ghosh and his team expressed their satisfaction with the accuracy and performance of the navigation system. | https://drive.google.com/open?id=12u3xCYITM20xI3PUqAW1ePrt9Mkymzve | Yes |  |  | NA | NA | NA | NA | The successful fusion of datasets and consistent accuracy highlight the robustness of the system. No additional concerns were raised during the case. | https://drive.google.com/open?id=1HvN8qj3JqpNLMhAjdqAVuJ7T5hR-g0yE, https://drive.google.com/open?id=1v97osi1JJ4JkOCanRptBvnHIEnMEpZl9, https://drive.google.com/open?id=1ZdXAzwL780A7hIz7qJSNwhYRiH99EwET, https://drive.google.com/open?id=1tSez8IbhWp5eZbnSk6dQmkSkgbDpf4J1, https://drive.google.com/open?id=1vq0OF5svlgaQHXhiW524uzhxilFR5ldt, https://drive.google.com/open?id=1_MqyjzrTM2JffQAOM_q0jqBjodfF_OOP, https://drive.google.com/open?id=1-9IoiOaQcmW-E_BxugqKvrxLE4nhDDMJ, https://drive.google.com/open?id=1UbYv3wI6T_n9gyphZ37DnUgZ2p8rgNUE | [12 empty] | Noted
2024-12-23 21:45 | rohan@rxoom.com | Rohan Ghosh | INK Kolkata | Kolkata | NS TRL 01 SE 101 | NS TRL 01 SE 101 | 1.02 | NA | https://drive.google.com/open?id=1Uf_gdXmi3RgNJLkF5hx3oHgdU5cnSNW3 | 2024-12-23 | Rohan Ghosh | Dr. Amit Kr Ghosh | Cranial |  |  | 0.1 | 4.52 | 0.56 | https://drive.google.com/open?id=19rSkuCp3ErjHhH4uohXk_kqS-Tr7UuZZ | Initial registration was completed, and the Junior Doctor observed that while accuracy was acceptable on fiducial points, there was a 1-2mm error along the midline. | The surgery concluded successfully, with Dr. Amit Kr Ghosh and his team expressing satisfaction with the overall accuracy and system performance. | https://drive.google.com/open?id=1RDWNTaBLcSFizM1wR1GuqgPZ1b4VZE7u | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1AL7dFDzQSu850NtHKjYG4E9pgOG_-gpm | https://drive.google.com/open?id=12tKPInNHp5E0n5xvdUfbkmFXmoRkExqZ | [11 empty] | Noted
2024-12-27 09:43 | trushalgulhane1547@gmail.com | Trushal Gulhane | Ciims hospital | Nagpur | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1GhrjtywamuHbkIcrwqcoEwM7QQhipt_C | 2024-12-25 | Trushal Gulhane | Dr. Vishal Dabare | Cranial |  |  | 0.3 | 3 | 0.9 | https://drive.google.com/open?id=14lEboNON9uYl59d5LuOSl9RVvBuS5cuE | Table adaptor block issue | Good |  | Yes |  |  | NA | Table adaptor block | NA | NA | NA | https://drive.google.com/open?id=1deJdu_u2AJCcCL2xjfdvZ_KtPm01b0HO | [12 empty] | Noted
2025 Demo Data
2025-01-03 22:45 | rohanghosh.2009@gmail.com | Rohan Ghosh | INK, Kolkata | Kolkata | NS TRL 01 SE 101 | NS TRL 01 SE 101 | 1.02 | NA | https://drive.google.com/open?id=1ZngB4Je8fncpnSkkM1cyHXOlZJJWsDG7 | 2025-01-03 | Rohan Ghosh | Dr. Amit Kr Ghosh | Cranial |  |  | 0.13 | 4.61 | 0.91 | https://drive.google.com/open?id=1sj8jH-GUi_zwxseB73wPqnjUMcj495Ir | NA | he navigation system demonstrated exceptional precision and accuracy, performing flawlessly with utmost reliability and precision always. | https://drive.google.com/open?id=1wUVvIG7CMis91zoid_eg3HzQ99a8F_QW | Yes |  |  | NA | NA | NA | NA | Today's craniotomy surgery at INK Hospital stands as a testament to the effectiveness of our navigation system. The successful integration of CT contrast data and the system's flawless performance highlight its capability to support complex procedures. This positive feedback from Dr. Amit Ghosh and his team reinforces our commitment to delivering advanced and dependable surgical solutions. | https://drive.google.com/open?id=144JNmQbpekJDNQqvwmV7Oq9vVH5ApSDA, https://drive.google.com/open?id=1QDoUzJc9zcDiOg0lprIhi31S5F-lx5pC | https://drive.google.com/open?id=1RFuN-QhX0o3PYxAcdaW7VO-2fNxXv-wg | [11 empty] | Noted
2025-01-04 18:26 | trushalgulhane1547@gmail.com | Trushal Gulhane | Ciims hospital nagpur | Nagpur | NS TRL 02 MF 201 |  | V1.2 | Yes 2 table block adaptor | https://drive.google.com/open?id=1Pr5dGnS9AlrBwtnWs_Qtbnp6uRW8j0uZ | 2025-01-04 | Trushal Gulhane | Dr. Pankaj sarda | Cranial |  |  | 0.2 | 6.2 | 1.1 | https://drive.google.com/open?id=1m2Eq3lBfksp6w9e-S7Ks5zx4tupoGy6m | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1_Dhh5lkxKJj8I_1WpuwSi8py8fGdgFrv | [12 empty] | Noted
2025-01-08 08:17 | trushalgulhane1547@gmail.com | Trushal Gulhane | P d hinduja | Mumbai | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1qG7VjUx7yQ6U0fxCeBF06lq6Hw2uTseD | 2025-01-07 | Trushal Gulhane | B k Mishra | Cranial |  |  | 0.2 | 3 | 1 | https://drive.google.com/open?id=1G5KNDigbcoPmeSWEOnHfpA6MejmjOEjW | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1LDQ-jBw1Z8ZRDR4jGikERladc1N5uUfa | [12 empty] | Noted
2025-01-08 14:30 | dhavalepatil1100@gmail.com | Manoj Dhawale | Dinanath Mangeshkar | Pune | Other | Dinanath Mangeshkar | 1.2 | NA | https://drive.google.com/open?id=1tw5X2V9Q6fcCwhGRlLHR7nv3ZAHA7g7T | 2025-01-07 | Manoj dhavale | Dr Tushar Borde | Biopsy, Any (Cranial, Full body) |  |  | 0.16 | 2 | 0.5 | https://drive.google.com/open?id=1rbj9ac2ofFH96jAaHs4YQvkZacSFTLlH | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1QNopnYpTXrAOCvgzeGewGEY-dOEjxgT8 | [12 empty] | Noted
2025-01-20 18:57 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | R1.2 | NA | https://drive.google.com/open?id=1_4jXaStiJQPy1YRrGlk5taohbjgODJne | 2025-01-08 | Gurjit Singh | Dr. Aditya Gupta and Dr. Neeraj | Cranial |  |  | 0.14 | 3.71 | 1 | https://drive.google.com/open?id=1N3e2Afypr6uO7WfQhFowM1xH-VxvTMn8 | NA | Navigation performance was good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1mtXgS2oKXfl36hyVWKI3KS97TE4jb0I9, https://drive.google.com/open?id=1Qd15DA0DXV_Y_nEfXKWb3pJmLCoiCUFR, https://drive.google.com/open?id=12JnLZUjTX91uDoxvf_-63k5RH6FHAJd4 | [12 empty] | Noted
2025-01-20 18:58 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=15uRHEgFqx3HXNH2qB_iT5Q2IjCwBkGM7 | 2025-01-10 | Gurjit Singh | Dr Pawan Goyal | Cranial |  |  | 0.12 | 3.23 | 0.54 | https://drive.google.com/open?id=1RvES2nzwfdbyfui5lv-f0fNoIa4_82Ue | Provided the patient MRI which was not according to navigation and provide wrong patient data. | Accuracy is good |  | Yes |  |  | NA | NA | NA | NA | Design good arm for camera and do provide multiple options like Medtronic like O arm base spine navigation, stereotactic navigation. | https://drive.google.com/open?id=1-oHibzeQZxT0B8AB1yH53mfrBo0o0y5P, https://drive.google.com/open?id=1rYQBGXy10BMhX03AykLerFl5tVBSbVPl | [12 empty] | Noted
2025-01-20 19:02 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1XB3qMDJt2KBO92figamKwHb0ML97ISLl | 2025-01-18 | Gurjit Singh | Dr. Aditya and Dr Sanjeev | Cranial |  |  | 0.11 | 4.32 | 0.71 | https://drive.google.com/open?id=1s-fRwN2g4ryJW3SkUHqAO_CKmiBd22kL | NA | Accuracy was good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1aWipbNY6drRzAYZGYw9NMcT9Xj5vW2dl, https://drive.google.com/open?id=17Xhgd9IMf_mgbfGtGLMQjAcdGt-VdFh7, https://drive.google.com/open?id=1dkk_5aQ7GqAWwRyNWRATc1tjsmscA59- | [12 empty] | Noted
2025-01-20 17:59 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1g7WEvN2hFRSI-v87YolaZVZIsZFupddW | 2024-01-20 | Gurjit Singh | Dr Aditya Gupta and Dr Sanjeev | Cranial |  |  | 0.26 | 4.04 | 0.88 | https://drive.google.com/open?id=1L-Ti0uLKZy2IfEhUzW8ag8uB3ZQICPxs | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1l5V2x3YC8za2hCkEc3WW5FDBGVSlSZvh, https://drive.google.com/open?id=1IFiCjow0l9grc6E8th2ibPsPlYNz-qEz, https://drive.google.com/open?id=1WIVUaa8PTMHFaBnHyGPe75Y1OhEu1vG4, https://drive.google.com/open?id=1jo_lHNOGSoM8jjdm-cnrNfD71wcS8jNb | [12 empty] | Noted
2025-01-20 18:55 | gurjit758.232@gmail.com | Gurjit Singh | Artemis Hospital | Gurgaon | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1ysCaOKmNw8ArlHjTDY89KcUgp2e_qG3s | 2025-01-20 | Gurjit Singh | Dr Sanjeev and Dr Pawan | Cranial |  |  | 0.11 | 3.51 | 0.43 | https://drive.google.com/open?id=1vl3ADhkLQJb6iO_66izSznpfhbAApSvC | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1DgiyCpFblP49NudU4_9zYdINoxym25IB, https://drive.google.com/open?id=1IxHHou94qMR85SaKfOZMnwMcmTXxUn43 | [12 empty] | Noted
2025-01-31 12:57 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Manipal hospital | Bengaluru | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=15m0ybpHTcwIb92wgC50ISChW8kVY5N8K | 2025-01-30 | Hemarshitha H S | Dr Sathwik Shetty | Cranial |  |  | 0.2 | 3.24 | 0.99 | https://drive.google.com/open?id=1nktHZc6YATE9DoqytEcfoAn8H5Aa-2kC | The reference was initially requested to be fixed on the Sugita headframe. During registration, challenges were observed as the arm and headframe obstructed point collection in the prone position. | 1. A modified reference arm is required that can be securely fixed on the Sugita headframe.This is essential as the headframe is frequently adjusted intraoperatively. 2.The surgeon requested an increased length for the long probe to enhance maneuverability | https://drive.google.com/open?id=11NvZu_5s5eIDWADj1f2i-JaIIW3DvKms | Yes |  |  | NA | NA | NA | The adaptor should be manufactured to fit securely onto the Sugita head frame. | A wider camera field of view is required to improve tracking efficiency and ease of use | https://drive.google.com/open?id=1XTU_W7-cqwyukpPBY6BEaT7JBwUBGjly | [12 empty] | Noted
2025-02-01 21:51 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Manipal hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1_iuV0a6Rl2J5-iYJIsZ-vo1hIDuRE417 | 2025-02-01 | Hemarshitha H S | Dr Sampath chandrarao | Cranial |  |  | 0.2 | 4.58 | 0.48 | https://drive.google.com/open?id=1AG9Q5s8reA-QqfaIcPiKF660YP95c9p7 | NA | Surgeon was happy with the system performance |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1LillujAyRst_MJA36bD9M5oUP-nQN2vH | [12 empty] | Noted
2025-02-05 13:53 | trushalgulhane1547@gmail.com | Suraj Garad | Noble hospital | Pune | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1hUQ44R8ZNAodJ1GuX2144b7cWLdQKISw | 2025-02-05 | Trushal Gulhane | Dr. Chandrashekhar and dr. Santosh jagtap | CT-Spine |  | 10 | 0.2 | 4 | 0.78 | https://drive.google.com/open?id=1XWIM5zoEM4tHMZxb_feU5z3ywkjNDOlk | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1Bo-yagrHAig8v0HNiaJlgyHDaiNScVq9 | [12 empty] | Noted
2025-02-05 16:40 | garadsuraj567@gmail.com | Suraj Garad | Hinduja hospital khar | Mumbai | RX - MT 0923 - 01 |  | V1. 2.24 | New Biopsy tracker | https://drive.google.com/open?id=1NTYWvKlwToop55cH7nxXoZUFAHWAVh-A | 2025-02-04 | Suraj garad | Dr. B K Mishra | Cranial |  |  | 0.1 | 4.18 | 1.18 | https://drive.google.com/open?id=1Yd16-EEPdIZTtsrogMx_dxVWL1mnIK4n | After the draping we got some accuracy error | NA |  | Yes |  |  |  | NA | NA | NA | NA | https://drive.google.com/open?id=19_6lsuZyXIHKu1KbU1v4UyDctMzbHpba | [12 empty] | Noted
2025-02-05 21:11 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Manipal hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1BeR-i7dht8ipl-ITqeKatVFFHTe2tZGh | 2025-02-05 | Hemarshitha H S | Dr Praveen M G | Cranial |  |  | 0.13 | 3.79 | 0.41 | https://drive.google.com/open?id=1JxCEiDhtfKs62XgFI681ln7bwzOq0_sF | NA | * The surgeon noted that, in comparison to other navigation systems, the registration process was simpler and more efficient. * He particularly appreciated the skull tracker reference attachment. |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1pVXm56iggwrClXiSbXcMVDlSnbRfezYo | [12 empty] | Noted
2025-02-07 14:49 | trushalgulhane1547@gmail.com | Suraj Garad | Noble hospital | Pune | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1JtZenpVMR0HudSrrbYKFgc5mPK_3ighl | 2025-02-07 | Trushal Gulhane | Dr. Chandrashekhar | Cranial |  |  | 0.07 | 7 | 1.07 | https://drive.google.com/open?id=1EUNOf1d1WKM_z9LnhhZhTgM0nkk9Z9Do | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1U1T-fI1U38iNEvMOoGLJr1B8YwqclEhF | [12 empty] | Noted
2025-02-11 21:44 | trushalgulhane1547@gmail.com | Trushal Gulhane | P d hinduja hospital | Mumbai | NS TRL 04 DB 101 |  | V1.2 | NA | https://drive.google.com/open?id=1AYbeT32vrCja5qrV-nAY-nBFrjwoSJmB | 2025-02-11 | Trushal Gulhane | Dr. B k Mishra | Cranial |  |  | 0.02 | 13 | 1 | https://drive.google.com/open?id=1xtLkcBnnJOcArHr32hgSPWwXSLsXRxmX | Skull clamp adaptor moved | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1YgHfuauy5ztlUqg-H8QMmJxcxz27O9vM | [12 empty] | Noted
2025-02-14 13:34 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Manipal hospital | Banglore | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=197KUqEpL2eLtWMM3p89LDTphmbj5dp86 | 2025-02-10 | Hemarshitha H S | Dr Bopanna | Cranial |  |  | 0.03 | 2.6 | 0.86 | https://drive.google.com/open?id=1QogCpYTKp0yHjLQjWitfurc0jekPo6BP | NA | Rigidity of IR camera arm/ limited mobility | https://drive.google.com/open?id=184ErrXwWbtZEBrTK_HCxc1RwuBWPgF1y | Yes |  |  | NA | NA | NA | NA | The camera arm should be able to move both vertically and horizontally to avoid obstruction when using the microscope and probe simultaneously | https://drive.google.com/open?id=1w94FY7gn15Qp5V_17qxb7gfII1gojmrS | https://drive.google.com/open?id=1HsLUc2OQ1ZEIknRImVJBMKLR1B8EEEjk | [11 empty] | Noted
2025-02-20 20:59 | trushalgulhane1547@gmail.com | Trushal Gulhane | P d hinduja | Mumbai khar road | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2 | NA | https://drive.google.com/open?id=18KqSSGF46JGYjFgyGUIZKjyCRHYgiMQE | 2025-02-18 | Trushal Gulhane | B k Mishra | Cranial |  |  | 0.2 | 3 | 0.5 | https://drive.google.com/open?id=1PZTk5r4H5UQ0p6D7l8EWfR_P26fdLWFe | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1EH8R-dToWegWqtg6NfOwYaeR21NVT4E9 | [12 empty] | Noted
2025-02-26 10:44 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Manipal Hospital | Bangalore | NS TRL 02 MF 204 |  | V.1.02 | NA | https://drive.google.com/open?id=1wFSicyB9kD5SezZfuoF1sPGGSsdL983C | 2025-02-17 | Hemarshitha HS | Dr. Sampath Chandrarao | Cranial |  |  | 0.03 | 2.6 | 0.86 | https://drive.google.com/open?id=1H2Q2d_v4mJ36i4Ygu20ZJKLtLZRQkUkk | Initially, registration was completed using the new skull tracker, which was attached using double tape. Pre-draping, the system displayed perfect accuracy. However, after draping, there was a noticeable difference in accuracy. The system showed substantial errors, with one instance even indicating the probe was in the air. I informed the surgeon that the reference might have shifted, but the surgeon pointed out that if the reference had truly moved, it would show an error every time. Despite the system displaying perfect accuracy initially, it failed to maintain consistency. After further evaluation, the surgeon confirmed that the reference remained in place and stable throughout the procedure. Ultimately, the surgeon decided not to use the system for this case. | After further evaluation, the surgeon confirmed that the reference remained in place and stable throughout the procedure. Ultimately, the surgeon decided not to use the system for this case. |  | No | Maybe Software or Mini plate Location Block Moved | Software | NA | NA | Trolley Bag Was heavy & Wheels was not proper. But later Bag Changed | Trolley is too Heavy | NA | https://drive.google.com/open?id=12hxbu7mnCTdGYUveuCbug07flkTqtqPQ | [12 empty] | Noted
2025-02-26 17:01 | trushalgulhane1547@gmail.com | Trushal Gulhane | Jupiter hospital | Thane | NS TRL 02 MF 201 |  | V1.2 | NA | https://drive.google.com/open?id=1pipPwUI8oB_e_3YkhqQJZ1IgC0_AxxK5 | 2025-02-26 | Trushal Gulhane | Dr. Priyank Patel | CT-Spine |  | 20 | 0.2 | 3 | 0.5 | https://drive.google.com/open?id=1AciO37yTtzTT2SgXDN7EdJwnvkZlJ36V | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1zccfEIASFfIsdB5F3vNT-NlHt2zUW25V | [12 empty] | Noted
2025-02-26 19:21 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Jubilee Mission Hospital | Thrissur | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1P3ECk8Ohr-t86FnIGTfH3FOPKW5r9pwN | 2025-02-26 | Hemarshitha H S | Dr Ramesh | Cranial |  |  | 0.17 | 3.81 | 0.5 | https://drive.google.com/open?id=1QqmPjCr_s19-fKRqovvbbHy1hlV4iDd8 | NA | Surgeon expressed satisfaction with the registration process and accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1v26tG6vekZOJi8VoNUxsIWkRaCRZckaO | [12 empty] | Noted
2025-03-03 15:50 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Jubilee Mission Hospital | Thrissur | NS TRL 02 MF 204 |  | V1.02 | NA | https://drive.google.com/open?id=1fOKhEWQ5MphW1RdWWlnRvwXQzlTcnrTn | 2025-02-27 | Hemarshitha HS | Dr. Kurian, Dr. Neil | Cranial |  |  | 0.33 | 3.78 | 0.37 | https://drive.google.com/open?id=1TuEn9A0mKwyzm6SBCD9R5xgHyhYsXmuc | Post-Draping Accuracy Check:Before opening the skull, the surgeon cross-checked accuracy and observed that the probe appeared approximately 1mm inside the skull on navigation.The surgeon understood the discrepancy and adjusted navigation usage accordingly. | 1.Dr. Neil : * Stated that this system has the easiest registration process compared to other systems. * Found it simple and user-friendly. 2.Dr. Kurian : * Acknowledged that the registration process is fast. * Appreciated the software features. * Expressed satisfaction with the system’s overall performance. * Mentioned that since this was a surface tumor case, he would like to test it in a deeper tumor case for further evaluation. |  | Yes |  |  | NA | NA | System is Heavy | NA | NA | https://drive.google.com/open?id=1LkxdXijsERBdzuyBPvdAWPzKK_ogLr-x | [12 empty] | Noted
2025-03-03 15:54 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Jubilee Mission Hospital | Thrissur | NS TRL 02 MF 204 |  | V01.02 | NA | https://drive.google.com/open?id=1LznRi6_FdOiwh2ksxvcfFDlE1mxaPo0Y | 2025-02-28 | Hemarshitha HS | Dr. Ramesh, Dr. Varun | Cranial |  |  | 0.26 | 5.42 | 0.67 | https://drive.google.com/open?id=18RCymFaGC9zGNM0EvdWq9rf0VSweOFZ1 | NA | * Dr. Varun : Expressed that the registration process is easy across all patient positions. * Dr. Ramesh : Was satisfied with the accuracy and overall system performance intraoperatively. |  | Yes |  |  | NA | NA | System is Heavy | NA | NA | https://drive.google.com/open?id=1G0obLY0CiQjGrr85DklnKbW0IENjqo20 | [12 empty] | Noted
2025-03-04 16:47 | trushalgulhane1547@gmail.com | Trushal Gulhane | Jupiter hospital | Thane | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2 | Na | https://drive.google.com/open?id=19SstN5fhIaRh64STUCWxpVBS2OolkFZW | 2025-03-03 | Trushal Gulhane | Dr. Priyank Patel | CT-Spine |  | 20 | 0.3 | 6 | 1 | https://drive.google.com/open?id=1k0O0-f7SHB_VEzCrjh-p06TGsB-n8cPh | NA | Good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1LQTNn51_KwDmu3Z-9RN7rq3LpWWAfK55 | [12 empty] | Noted
2025-03-07 21:42 | gurjit758.232@gmail.com | Gurjit Singh | ERA hospital | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1YXj4up3Hbt7SC1fJy0TY41oV6xK1tvqz | 2025-03-07 | Gurjit Singh | Dr. Alok Srivastava | Cranial |  |  | 100 | 2.95 | 0.59 | https://drive.google.com/open?id=14x0--dJBnWVcdjLuxK5ktfjJ5LVHCVNB | NA | Feel Ease of user interface and workflows. |  | Yes |  |  | NA | NA | Heaviness of system. | NA | NA | https://drive.google.com/open?id=1VweoQgWfDUKavFQP5uG0bJweEdwCScfj | https://drive.google.com/open?id=1hbR32xIwCMomVLRYHN-QW7uKIhn-1Pyu | [11 empty] | Noted
2025-03-07 23:45 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Jubilee mission hospital | Thrissur | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=1FNhPfLSq2oAldiYzq317JeWT8r59mrV4 | 2025-03-05 | Hemarshitha H S | Dr Ramesh | Cranial |  |  | 0.45 | 6.76 | 0.82 | https://drive.google.com/open?id=1hxuFoS_FzxW9E8MYubOwRVUrPPQ9XlnZ | *The reference arm was blocked on both sides, preventing movement in the ball-and-socket region. * After completing the registration process, the 3D model froze. | Dr. Ramesh was highly impressed with the system’s performance. He expressed that he never expected an Indian product to compete with Medtronic at this level. He also assured that he would recommend the system for purchase to the hospital management |  | Yes |  |  | NA | NA | NA | Replacement of new reference arm | NA | https://drive.google.com/open?id=1t9XlRAtufJxvc32yFei2lcPviZ2FUYoK | [12 empty] | Noted
2025-03-11 17:10 | trushalgulhane1547@gmail.com | Trushal Gulhane | Narayana hospital | Ahemadabad | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2 | NA | https://drive.google.com/open?id=1D33DXoSMyn8yYReAuolBOHEC6lWsJMTX | 2025-03-11 | Trushal Gulhane | Nikunj Godani | CT-Spine |  | 20 | 0.3 | 7 | 1 | https://drive.google.com/open?id=1Sdno05ru4F-KzlB7eEfQH09ygD5ZB1ML | NA | Good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1g6G34TmlyUcitHsA1udWrHGMUupG54LS | [12 empty] | Noted
2025-03-11 23:11 | gurjit758.232@gmail.com | Gurjit Singh | ERA College and Hospital | Lucknow | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1yaQlPoBwJFPcymuKpTDdDP3EPPey2p8G | 2025-03-08 | Gurjit Singh | Dr Amir | CT-Spine |  | 30 | 0.38 | 1.94 | 1.17 | https://drive.google.com/open?id=1VkZ49xtgxqZ8P1SGN04N-jr7yn4hKxLY | NA | Good accuracy in spine but not much value addition in spine |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1GqiqrNIDPLosp62lmEjf3g2ilvYAqXNZ | [12 empty] | Noted
2025-03-12 19:17 | rohan@rxoom.com | Hemarshitha H. S. | JUBILEE MISSION MEDICAL COLLEGE & RESEARCH INSTITUTE | Thrissur | NS TRL 02 MF 204 |  | 1.2 | NA | https://drive.google.com/open?id=1y8T5-mAkLsW9G1GQikA9AoCkhTM23SVT | 2025-03-08 | Rohan Ghosh | Dr. Kurian & Dr. Vipin | Cranial |  |  | 0.17 | 5.88 | 0.5 | https://drive.google.com/open?id=1hjU7ExcILs6_djjuJtRl0osAiYC5NNud | - Towards the end of the procedure, the *camera took slightly longer than usual* to recognize the fiducials (Glaons) attached to the probes. - Upon investigation, it was observed that one of the three fiducials was obstructed—either by the surgeon’s fingers or another obstacle between the camera and fiducials. - The issue was promptly identified and corrected by the surgical team. | The navigation system performed efficiently during the procedure, with *high accuracy and seamless registration. The **positive feedback from the surgical team* reinforces the reliability of our system. Implementing additional software features, as suggested by the doctors, could further enhance its capabilities and competitiveness. | https://drive.google.com/open?id=1qJrzOkhDF2w3TsH8-TpR0H21IgDhpDa7 | Yes |  |  | The navigation system performed efficiently during the procedure, with *high accuracy and seamless registration. The positive feedback from the surgical team reinforces the reliability of our system. Implementing additional software features, as suggested by the doctors, could further enhance its capabilities and competitiveness. | NA | NA | Implementing additional software features, as suggested by the doctors, could further enhance its capabilities and competitiveness. | Implementing additional software features, as suggested by the doctors, could further enhance its capabilities and competitiveness. | https://drive.google.com/open?id=1SnUQrSBoJLqSFJ44tV_VS4LxtBR9zbcM, https://drive.google.com/open?id=1uMXTfYPQjWn16GZuWgt8AWEc9Jk0jEIf, https://drive.google.com/open?id=1Ypiqsd3_Aw0FZiQ3IEk_OHslRriQZdTH | https://drive.google.com/open?id=1aV9_kZnwgSENxFZpOQqU_Ix1_mGjLoKo | [11 empty] | Noted
2025-03-15 11:13 | trushalgulhane1547@gmail.com | Trushal Gulhane | Epic hospital | Ahemdabad | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2 | NA | https://drive.google.com/open?id=1bxakkOROerAOPNwULKFgdWZmYeY_cJAo | 2025-03-12 | Trushal Gulhane | Dr. Ankita and dr diwanji | Cranial |  |  | 0.2 | 3 | 0.6 | https://drive.google.com/open?id=1J__txHuavDqNksYnIS6fjDi7yRgMQ6EL | NA | Good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=12ytQZTN0YiA1aFQI2ejpiolG8a_WhS09 | [12 empty] | Noted
2025-03-15 11:17 | trushalgulhane1547@gmail.com | Trushal Gulhane | Narayana hospital | Ahmedabad | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2 | NA | https://drive.google.com/open?id=1Sb7uGhuT_QM-5u2uNRgi90CVLWy5HksO | 2025-03-13 | Trushal Gulhane | Dr. Nikunj Godani | Cranial |  |  | 0.3 | 2 | 0.7 | https://drive.google.com/open?id=1qYRb0V2Xy8F5ZjhFQK9raZGK-Sx4A4Wh | NA | Good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1E_XPB_wora_6QMaq1W5AHKKmMDtK3u4q | [12 empty] | Noted
2025-03-17 22:15 | trushalgulhane1547@gmail.com | Trushal Gulhane | Narayana hospital | Ahemdabad | NS TRL 04 DB 101 | Middle pipe hexagonal | V1.2 | NA | https://drive.google.com/open?id=1zwuUUR8trFM4gQTF0JxtZN10jOx2FKWA | 2025-03-17 | Trushal Gulhane | Nikunj Godani | Cranial |  |  | 0.3 | 4 | 0.5 | https://drive.google.com/open?id=193eEs35dHN-m_CpxML8hPzl1CGhkJEew | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=17BFhJ4ar8EGZDHQhk4HPaqGpfUnxgkY2 | [12 empty] | Noted
2025-03-18 18:27 | trushalgulhane1547@gmail.com | Trushal Gulhane | Narayana hospital | Ahemdabad | NS TRL 04 DB 101 | Middle hexagonal shape pipe | V1.2 | NA | https://drive.google.com/open?id=1fBf9a75oZUavLNASHPFk0xFQdHYiHGV- | 2025-03-18 | Trushal Gulhane | Dr. Nikunj Godani | Biopsy, Any (Cranial, Full body) |  |  | 0.2 | 4 | 0.4 | https://drive.google.com/open?id=1C8C0KEU50rZbWo3jeOrVffO5vRm9aZyQ | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1pqj3LylOYcjj_PwzkUzf0qLwqRpz5Lbz | [12 empty] | Noted
2025-03-18 22:22 | gurjit758.232@gmail.com | Gurjit Singh | Graphic ERA Hospital | Dehradun | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1N0x8ejvqaxDpqCy3mF9YT65uktLiYOYP | 2025-03-18 | Gurjit Singh | Ankur Kapoor | CT-Spine |  | 10 | 0.35 | 3.37 | 0.46 | https://drive.google.com/open?id=1_22tSBlrJ3WSaC7FB-5_b68gCf96JVQD | NA | Good Accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=17SpGvsEAA_mSX6UNTK28SwW3n1phc6y0 | https://drive.google.com/open?id=1m6n4204pTS0dCEsokoW6m9qiryatANp5 | [11 empty] | Noted
2025-03-22 22:35 | gurjit758.232@gmail.com | Gurjit Singh | Graphic ERA Hospital | Dehradun | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1ioPbT1en7wdQGdedJtI7yZPhYkvuEJ6W | 2025-03-22 | Gurjit Singh | Ankur Kapoor | Cranial |  |  | 0.12 | 4.82 | 1.88 | https://drive.google.com/open?id=1XRpENCw2JFquUfEprq_c_ZDN505k7Gd5 | Felling short of screws | Have better way of fixation of skull tracker |  | Yes |  |  | NA | One passive marker | NA | NA | NA | https://drive.google.com/open?id=1aDEmpyay_9CYvF18XV5XcDdRr8RAcnrp | [12 empty] | Noted
2025-03-23 23:21 | hemarshitha1802@gmail.com | Hemarshitha H. S. | Jubilee mission hospital | Thrissur | NS TRL 02 MF 204 |  | 1.02 | NA | https://drive.google.com/open?id=17bAYrJhX38JWVXhjtqe7mj4g97W2RcGw | 2025-03-13 | Hemarshitha H S | Dr Kurian | Cranial |  |  | 0.22 | 4.79 | 0.42 | https://drive.google.com/open?id=1XIhOH8vvxB6y0rDrVDVvnfE6f1XUnMxF | NA | ✔ Mentioned that the registration process was simple when accuracy was good. ✔ Found the overall system user-friendly and efficient. |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1V4lZmvPJ7AnVkOlbyg-9JsH1ess5M1bb | [12 empty] | Noted
2025-03-25 17:24 | gurjit758.232@gmail.com | Gurjit Singh | Graphic ERA hospital | Dehradune | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=140n23GG5X-gbef_i155VL4t4C6_wHeLr | 2025-03-25 | Gurjit Singh | Dr Ankur Kapoor | X-Ray Spine |  | 10 | 0.35 | 100 | 1.18 | https://drive.google.com/open?id=1U2dQo1JqWvoy73XMuX3OLhgJRQlNQBxA | NA | Good accuracy, useful |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1g-WqjxHBOoCxmp0JoH7HekkPr1PrE57z | [12 empty] | Noted
2025-03-26 17:51 | gurjit758.232@gmail.com | Gurjit Singh | Graphic Hospital | Dehradune | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1G36KT_nuQjrxgWGr1RFsloe45HWDG36u | 2025-03-26 | Gurjit Singh | Dr Ankur | Cranial |  |  | 0.2 | 3.62 | 0.59 | https://drive.google.com/open?id=1NLH_drzxWJhHMHUbNHldtMdt6a8Ubx3u | NA | Good Accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1ndecZiqrDxTybOHUbfHGkE2AIMtpEdjR | [12 empty] | Noted
2025-03-28 20:17 | gurjit758.232@gmail.com | Gurjit Singh | Graphic ERA Hospital | Dehradune | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1kF0RmOR1DrrUm5KLOV5hOQISca4-cOGs | 2025-03-28 | Gurjit Singh | Dr Ankur Kapoor | Cranial |  |  | 0.12 | 1.33 | 0.52 | https://drive.google.com/open?id=1pQ-G2SEIdPLd9D7DvrniSg9XmBWuEfWN | NA | Good Accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1B8jbxyQ1dK2zZvH2uXG5yx0HTmeIwQqT | [12 empty] | Noted
2025-03-31 23:36 | rohan@rxoom.com | Hyderabad Distributor | Apollo Hospital, Secunderabad Branch | Secunderabad | NS TRL 01 SE 101 | NS TRL 01 SE 101 | V1.2.24 | Pedicel Finder | https://drive.google.com/open?id=1Dq1IOEuDAwbR3ek7OPXukSrBQKVpKBVb | 2025-03-31 | Rohan Ghosh | Dr. bramha Prasad | CT-Spine |  | 15 | 0.09 | 2.87 | 0.58 | https://drive.google.com/open?id=17wLFxWJPFCiSTsNhVPD7DZmB73lVFG_M | 1.The hospital provided a CT dataset of approximately 3GB, containing over 1500+ slices. This dataset was four times larger than standard CT data, causing the application to hang upon upload. 2. The model creation process during navigation took significantly longer due to the large data size. 3. The application froze multiple times during fiducial registration, leading to delays in accurately marking the points. 4. The doctor requested the integration of a real-time screw-fixing software sequence for enhanced procedural efficiency. | 1.Ease of Loading Patient Data: Agreed 2. Registration Process Simplicity: Strongly Agreed 3. Navigation Accuracy: Agreed 4. Ergonomic Design of Instruments: Strongly Agreed 5. User Interface (UI) Simplicity: Strongly Agreed 6. Additional Comments & Suggestions: Additional fiducials on the background and an overlay feature for the end of pedicle screws would be helpful. | https://drive.google.com/open?id=11Aloyi7pBoszzSi1xKWwmw9FijYZx5FX | Yes |  |  | Additional fiducials on the background and an overlay feature for the end of pedicle screws would be helpful. | NA | NA | 1.Additional fiducials on the background and an overlay feature for the end of pedicle screws would be helpful. 2. We require the design and development of a robust application capable of efficiently handling large-sized data sets. | Additional fiducials on the background and an overlay feature for the end of pedicle screws would be helpful. | https://drive.google.com/open?id=1Aabxew03g8edhciikbu9bF034P2pumG3 | https://drive.google.com/open?id=1xqohOD8WeX4XfSp_oXcB4zJaZVsYniVv | [11 empty] | Noted
2025-04-02 21:24 | gurjit758.232@gmail.com | Gurjit Singh | Max Super Speciality Hospital | Dehradun | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1LzIDBk0FcNcAhjNiii-ZeXeiCuX3nZMh | 2025-04-02 | Gurjit Singh | Dr. Priyank Uniyal | CT-Spine |  | 30 | 0.27 | 4.35 | 1.38 | https://drive.google.com/open?id=1dvCdGtb8L97NghryrtDGqEzz8FG6Kq66 | NA | Good accuracy, easy registration |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1qt12AAOyHGA7f8lgo9_lfxLNK6Bz2q9c, https://drive.google.com/open?id=18TB2H2M6_FJDvX7Z58o4stAICcQ2K75D, https://drive.google.com/open?id=1eNVaC6NDDWgOI_9mfTDgyx1EWYb6V1vi | https://drive.google.com/open?id=1SmtSyCyHMxu5dRY2VdHDdWW7dejRGdF9 | [11 empty] | Noted
2025-04-06 | gurjit758.232@gmail.com | Gurjit Singh | Max Hospital | Dehradun | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1ELcMmY5HaVc4MOtLPiCfM-oA22wnUhLv | 2025-04-05 | Gurjit Singh | Dr. Mukesh | Cranial |  |  | 0.22 | 2.42 | 0.54 | https://drive.google.com/open?id=1BFnbMnUt9As3UpA6F3e4REK15T2j3Ytl | NA | Good accuracy, Easy to use |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1UYILrjDg59ALUMcLfsAhI5BZsTm0c08i, https://drive.google.com/open?id=1tE9S_fIkTqoBcMa00KMgLQanGc7kTXVN | [12 empty] | Noted
2025-04-06 22:04 | trushalgulhane1547@gmail.com | Trushal Gulhane | Hinduja hospital | Mumbai | NS TRL 04 DB 101 | Middle hexagonal shape pipe | V1.2.24 | Na | https://drive.google.com/open?id=1-O78WI83wpjvKRyEvomLCBk6E3ZFd-9O | 2025-04-06 | Trushal Gulhane | Dr b. K. Mishra | Cranial |  |  | 0.2 | 2.5 | 0.4 | https://drive.google.com/open?id=1MD-TA3AuEHn_cL_P7w2j9TTi9eBZ6O6o | Model got disorder while during the registration | NA |  | Yes |  |  | Na | Na | Na | Na | Na | https://drive.google.com/open?id=1Pzo-vZf_InJBNvvIWarEYtAADTFWNd-X | [12 empty] | Noted
2025-04-12 20:56 | trushalgulhane1547@gmail.com | Trushal Gulhane | Hinduja hospital | Mumbai | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=1eEZFlzE8Ei6Fprfj7vEzwSoyd_EtiyPR | 2025-04-08 | Trushal Gulhane | Dr b k Mishra | Cranial |  |  | 0.2 | 2 | 0.5 | https://drive.google.com/open?id=14aYzctcEl7jWn_J1cLwP_UpZN5PSfe44 | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1wYUoOF7LDD_iKoCRL64Bl6XUh7NkOrWB | [12 empty] | Noted
2025-04-12 21:01 | trushalgulhane1547@gmail.com | Trushal Gulhane | Hinduja hospital | Mumbai | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=1RpbeHAorpWTbNMT9mYAXpVaPMr_X8O4m | 2025-04-12 | Trushal Gulhane | Dr anshu warade | Cranial |  |  | 0.2 | 5 | 1.2 | https://drive.google.com/open?id=1gc0JdGHeRpLgQ5m3Q97ICA6p8WPZh306 | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1ypT7to2j2RxeFVROBU9uTHaQRywaAkdD | [12 empty] | Noted
2025-04-15 20:51 | gurjit758.232@gmail.com | Gurjit Singh | MAX hospital | Noida | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1Awb8X4Toy-xcw1vPWa6IQ6qnZkLeqaLQ | 2025-04-15 | Gurjit Singh | Dr. Rohan Sinha and Dr. Dinesh Rattnani | Cranial |  |  | 100 | 3.88 | 0.43 | https://drive.google.com/open?id=1ShqSXx2g_2-gUmVhIzYhc_XL5MbNmCID | NA | Good accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1-FMMHdONzJuk7ZsgD4--wuRb0lBhkO8r | [12 empty] | Noted
2025-04-18 21:09 | trushalgulhane1547@gmail.com | Trushal Gulhane | Hinduj hospital | Mumbai | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=12BFwVizVPVWotd7xhEl_mVHiKlhnkDVE | 2025-04-18 | Trushal Gulhane | Dr b k Mishra | Cranial |  |  | 0.2 | 6 | 1.2 | https://drive.google.com/open?id=1q7c9uvbA654yaqPybc-saJvLAniZI52L | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1IBwPREbqBzQmnAmKnStsxofZaMl1Xa67 | [12 empty] | Noted
2025-04-19 14:52 | dhavalepatil1100@gmail.com | Manoj Dhawale | Nobal hospital | Pune | NS TRL 02 MF 206 | Noble Hospital system | V1.2.24 | NA | https://drive.google.com/open?id=1COMH7fHzO128S_CVuI5JL5J0Rt5cRyb5 | 2025-04-12 | Manoj dhavale | Chandrashekhar raman | Cranial |  |  | 0.1 | 4.34 | 0.65 | https://drive.google.com/open?id=1PQL-TSVjEYnYemcv0zTzuH_Zl2R_3PbK | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1DBYqRLvXeYIv5RaZfkGXH5jgz5mie2l2 | [12 empty] | Noted
2025-04-25 16:50 | gurjit758.232@gmail.com | Gurjit Singh | Max Hospital | Noida | NS TRL 02 MF 201 |  | V1.2.24 | NA | https://drive.google.com/open?id=1WPrTCdJOsaXdWK6KUHrFhephBccYqH2t | 2025-04-25 | Gurjit Singh | Dr. Rohan Sinha | CT-Spine |  | 20 | 0.72 | 2.53 | 0.41 | https://drive.google.com/open?id=1zm_meQSA-z8raOZ7oveDyVQxkhp2tGvt | NA | Good accuracy on register vertebrae but not on others and says they need to get used to for registration for different vertebrae. |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1i2vJF-x39WLGxj46cAtj6vEv8vzk3lzB | [12 empty] | Noted
2025-04-26 10:33 | hemarshitha1802@gmail.com | Hemarshitha H. S. | BMH Hospital | Kannur (Kerala) | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1B4989tGH8nvOBa2B63Kvz1giNe8PSz7h | 2025-04-24 | Hemarshitha HS | Dr. K T Suhas | Cranial |  |  | 0.72 | 4.04 | 0.38 | https://drive.google.com/open?id=1YD_RhqXL-Ykh1asesABjzhx_BDRTK4it | NA | Dr. Suhas reported that the system was user-friendly, maintained excellent accuracy throughout the procedure, and the instruments were comfortable and easy to handle. | https://drive.google.com/open?id=1ewA7R4cl3kigsz9D-buvjKDbU1FALxWA | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1Va1bVQHAr8QDM3jgW6qYGuWuHkhR1ID2 | [12 empty] | Noted
2025-04-26 10:38 | hemarshitha1802@gmail.com | Hemarshitha H. S. | BMH Hospital | Kannur (Kerala) | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1pgd0ac8tuzuphqbT2pT1xDJ18yOVrGD6 | 2025-04-25 | Hemarshitha HS | Dr. K. T. Suhas | Cranial |  |  | 0.21 | 2.29 | 0.84 | https://drive.google.com/open?id=1nr6N1fpDLDE3aX9TbNdpWiaFudM0wRFc | NA | Dr. Suhas expressed high satisfaction with the registration technique, describing it as unique and effective. He found the system user-friendly and easy to handle, and he appreciated the excellent accuracy maintained throughout the surgery. | https://drive.google.com/open?id=1opAlwTQXz17mdOQxQk3ZmWGnaLCbZqMY | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1Qt1Qs5YnNVPJZNGXOsTPS0U8bkruxYv7 | [12 empty] | Noted
2025-04-26 22:04 | trushalgulhane1547@gmail.com | Trushal Gulhane | Sahyandri hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=1uZciDYmp7zJJyFeXuwEzvD8tWGNyE3Ie | 2025-04-26 | Trushal Gulhane | Dr. Rahul modgi | Cranial |  |  | 0.2 | 4 | 0.9 | https://drive.google.com/open?id=1VcTUpEnKNiqkQtxFG9Fb2QkgfqA2T__w | NA | Good |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1RToObYCgL0_cfyRBmCxalPOKrOvXA_4J | [12 empty] | Noted
2025-04-29 22:09 | trushalgulhane1547@gmail.com | Trushal Gulhane | Medicover hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=154bGuFTDe2NiBFTu1MfpHEXQnGFcoWIM | 2025-04-29 | Trushal Gulhane | Dr. Chrimade | Cranial |  |  | 0.1 | 1.8 | 0.9 | https://drive.google.com/open?id=1_92Tx1rHTZyy4ssxIbH2DfHFmnHKjeZO | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1V5DfzDA3PxfHruNnkm4Xp2S01v9V22bv | [12 empty] | Noted
2025-05-02 20:43 | gurjit758.232@gmail.com | Gurjit Singh | Max Hospital | Noida | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1n2FMR3WKpW2pvH6FDeTcUuiJ16m5q0IX | 2025-05-02 | Gurjit Singh | Dr. Rohan Sinha | Cranial |  |  | 0.26 | 1.55 | 1.05 | https://drive.google.com/open?id=19Kt_8hVBB20KN2SM6H7oTxtamJ_A4zZq | NA | Good accuracy and helpful system |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1hn3S0CoHa_bJlT2btI2MltyWhwFtm8t7 | [12 empty] | Noted
2025-05-03 13:40 | garadsuraj567@gmail.com | Faiz Khan | Noble hospital | Pune | NS TRL 02 MF 206 | NS TRL 02 MF 206 | V1. 2.24 | NA | https://drive.google.com/open?id=1vFNZ0j5vDQ5DbgwKgn-_0B_aanUKmerC | 2025-04-28 | Suraj garad | Dr. Chandrashekhar Raman | Cranial |  |  | 0.09 | 6.05 | 0.72 | https://drive.google.com/open?id=1wrRcx_nCiLy5xfLwe21348KWz4yhzSmB | Radiology department give us different data set which merge we have to take new cd from them for surgery | Na |  | Yes |  |  |  | NA | NA | NA | NA | https://drive.google.com/open?id=1o5JoPRhq7FudVxt8Lf4-3v_9CPKEQRU4 | [12 empty] | Noted
2025-05-03 13:48 | garadsuraj567@gmail.com | Faiz Khan | Noble hospital | Pune | NS TRL 02 MF 206 | NS TRL 02 MF 206 | V1. 2.24 | NA | https://drive.google.com/open?id=1xVuoeMSyVtzROK7e7ke2luqs5xGK6UsI | 2025-04-30 | Suraj and manoj | Dr. Chandrashekhar Raman | Biopsy, Any (Cranial, Full body) |  |  | 0.13 | 5.89 | 0.6 | https://drive.google.com/open?id=1kXzRfrwt__0xKM3kE12x6vhHT5Uti2PA | Na | Na |  | Yes |  |  |  | Na | Na | NA | Na | https://drive.google.com/open?id=1hlYofBrTNDiOddIzmxj3W3P_5O73xne_ | [12 empty] | Noted
2025-05-03 23:14 | rohan@rxoom.com | Rohan Ghosh | Tx Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=1yejF6yYRY_OaGAbOtCQy-WuNv3MOV3Fa | 2025-05-03 | Rohan Ghosh | D2 Spine tumor | CT-Spine |  | 0 | 0.07 | 3.66 | 0.83 | https://drive.google.com/open?id=1Wa9E0_Z-HyiJLy1XGq6DLXRPbhsQJ8lm | NA | Ease of loading patient data: Strongly Agreed Registration process simplicity: Strongly Agreed Navigation accuracy: Marked as Neutral (room for improvement) Ergonomics of instruments: Agreed Simplicity of User Interface (UI): Strongly Agreed He emphasized interest in conducting additional cases, specifically cranial surgeries and X-ray-based navigation, before considering further steps. | https://drive.google.com/open?id=14lgQ7AEStKPLlm7c7q_fAaV2vcGbbeAC | Yes |  |  | NA | Key board loss | NA | NA | Doctor emphasized interest in conducting additional cases, specifically cranial surgeries and X-ray-based navigation, before considering further steps. | https://drive.google.com/open?id=1-kONEJym_0gMzuidjxoOMqU3RX18T1De | https://drive.google.com/open?id=1jnjbaq1pFg5DjVIl57B8YHJmvQv-h3OF | [11 empty] | Noted
2025-05-07 23:02 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=13J64UXn8sZJt3FYgo-OHKZTinKm8WrlG | 2025-05-07 | Rohan Ghosh | Dr. TN Rao & Dr. Manick Prabhu | CT-Spine |  | 0 | 0.1 | 3.65 | 0.75 | https://drive.google.com/open?id=1u0l72zgePmWgoA4UL23w9nivPSFr2hfv | During screw fixation, minor vertebral movement resulted in a positional deviation, leading to a registration accuracy variance of approximately 1–2 mm. | During screw fixation, a minor positional shift in the vertebrae introduced a 1–2 mm deviation. However, the doctors confirmed that the navigation accuracy remained within clinically acceptable limits. | https://drive.google.com/open?id=18Aa8Bl_6RNASNdjkd1M-JUQ4dlZUa3JP | Yes |  |  | NA | NA | NA | NA | Based on the positive experience, Dr. T.N. Rao and his team expressed interest in conducting 3–4 additional surgeries to further evaluate and validate the system’s performance. | https://drive.google.com/open?id=1Ljeht31uaZCCFcbw7l7q70mT3khg3Gju | https://drive.google.com/open?id=17jZLiCSH3VklF4ydr2jlCCkZFFMOsGzr | [11 empty] | Noted
2025-05-08 10:15 | garadsuraj567@gmail.com | Faiz Khan | Dinanath hospital | Pune | NS TRL 01 MF 101 |  | 1.2.24 | NO | https://drive.google.com/open?id=1HhGLVM0oAf0TBmRspEmrrUk69743snh8 | 2025-05-07 | Suraj Garad | Dr. Nilange | Biopsy, Any (Cranial, Full body) |  |  | 0.05 | 5.77 | 0.3 | https://drive.google.com/open?id=1EOlgp4yqDTZQHBYJpPm1ZhMEv2s-6_9W | NO | NO |  | Yes |  |  | No | No | When we are shifting the trolley from one ot to another ot trolley is not moving easily we have to push with extra force. | NO | NO | https://drive.google.com/open?id=1qONxfMNm0-IxAchHVaC2nN3zDbyaLRMC | [12 empty] | Noted
2025-05-08 10:18 | garadsuraj567@gmail.com | Faiz Khan | Dinanath hospital | Pune | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1p0Z2c_M8bcsExafJ3BLR-lv5DOaAUZ4I | 2025-05-07 | Suraj Garad | Dr. Akshay Gupte | Cranial |  |  | 0.19 | 10.79 | 0.8 | https://drive.google.com/open?id=10doEqCTwr5RDP9UVOvWafrTJ7KhtZXIv | NA | NA |  | Yes |  |  | Na | NA | NA | NA | NA | https://drive.google.com/open?id=1sXLxExnX0mJdSWRC5Rxj2CxJYruG_p-t | [12 empty] | Noted
2025-05-08 23:37 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=13bXuytnIas1YerYuifE5Di_jEKAiYbgh | 2025-05-08 | Rohan Ghosh | Dr. Venkatesh Yadulla, Dr. Bhubaneswara R | Cranial |  |  | 0.23 | 4.25 | 1.43 | https://drive.google.com/open?id=16dtB0xaQI03xJXprEzmKxkfc8xrrkwu7 | Fiducial markers were placed, and the initial registration was performed by Dr. Bhubaneswara. While the system reported high fiducial registration accuracy, the visual alignment on screen suggested a mismatch—the probe appeared offset from the skin surface. | Dr. V. Yadulla and Dr. Bhubaneswara R expressed high satisfaction with the NeuroSaphireX system during the redo right CP angle tumor excision at Care Hospital, Banjara Hills. They strongly agreed that the patient data loading was easy, the registration process was simple, navigation was accurate, instruments were ergonomic, and the user interface was intuitive. The reported surface registration accuracy was 1.43 mm. The team appreciated the system’s reliability and performance throughout the surgery, indicating confidence in its capabilities and workflow efficiency. | https://drive.google.com/open?id=1XclQW2tQV80qLD1HfOXKHV5kFRBQB5TZ | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1BETiXk7F5M5w-22mxIOMQNFYo1WRapQI, https://drive.google.com/open?id=1dclksvaXjUBBj_nm7IUw-YfVuoDDSSHg | https://drive.google.com/open?id=1CyIruxuGa4VZkDrF2sF01BXMwTWmQNXT | [11 empty] | Noted
2025-05-12 23:52 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | Skull Tracker | https://drive.google.com/open?id=1zz8NUveT4Y8TvWZM-7HjjkF6Cfdxn83b | 2025-05-12 | Rohan Ghosh | Dr. TN Rao & Dr. Manik Prabhu | Cranial |  |  | 0.08 | 2.32 | 0.91 | https://drive.google.com/open?id=1xflQhcTvTTgqQliFa4XwnNWS5-lmmD_c | the navigation system failed to detect the long probe correctly. After thorough troubleshooting, it was discovered that the passive markers on the probe had been inadvertently replaced with older markers from the hospital’s Medtronic S7 system. Upon identifying this, the correct SaphireX passive markers were reattached, and normal functionality resumed with no further issues. | In a follow-up discussion, Dr. T. N. Rao expressed satisfaction with the NeuroSaphireX system, particularly noting: 1. High registration accuracy. 2. Simplicity in loading patient data. 3. Intuitive and efficient registration process. 4. Effective performance of the Skull Tracker in frameless navigation scenarios. | https://drive.google.com/open?id=1RsvPp9OvV-tzpot-MJUqUM6_1ojq5A5H | Yes |  |  | NA | NA | NA | NA | While the outcome of this case did not require tumour excision, the surgical team showed interest in evaluating the system further through more complex and varied neurosurgical cases to thoroughly assess its long-term performance and reliability. | https://drive.google.com/open?id=13ecR6IdRag4Iq-J860L4hBqfejNjSgbH, https://drive.google.com/open?id=1vSm1pD_jlU0NXM_1cpKDHmLlKuCXldMV, https://drive.google.com/open?id=1w9Xu9CL2WhTaV4qTO-o51FDHBTOWUfKa, https://drive.google.com/open?id=1-sXa__fNLgUUkEA2VX0MGXGp9k2wY3P2 | https://drive.google.com/open?id=1z3iDKiCJRfl68HODma5wu6vpG1x6oPMV | [11 empty] | Noted
2025-05-13 09:33 | trushalgulhane1547@gmail.com | Trushal Gulhane | Kims manavta hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | 2.2.24 | NA | https://drive.google.com/open?id=1K7HvNckVqymIPPHFibJU5lcx6Ey8Qn3L | 2025-05-10 | Trushal Gulhane | Dr. Ishan nerkar | CT-Spine |  | 20 | 0.17 | 2 | 0.45 | https://drive.google.com/open?id=17tEZG5e8SnN813uXZzT5N5EH1INQb18C | Play and movement in clamp | NA |  | No | NA | Big of the clamp for cervical spine fixation | Na | Na | Na | Please design small size clamp | Same as per above | https://drive.google.com/open?id=16fxcP4DC8CbQv5fuAD1CzFUr7U5ZIPj5 | [12 empty] | Noted
2025-05-13 14:15 | garadsuraj567@gmail.com | Faiz Khan | Noble hospital | Pune | NS TRL 02 MF 206 | NS TRL 02 MF 206 | V1. 2.24 | NA | https://drive.google.com/open?id=1pkHaPTGd9KEpL332jUxKLXCB5-oHjlpo | 2025-05-13 | Suraj garad | Dr. Chandrashekhar Raman | Cranial |  |  | 0.05 | 6.14 | 0.78 | https://drive.google.com/open?id=1CdA0jPf9p1uVNVwXOExIo1SAa5M8vNjY | No | NA |  | Yes |  |  |  | NA | NA | NA | NA | https://drive.google.com/open?id=1qTTh7TNOyRvmFBh4fRTb9k8cs4KhMjg- | [12 empty] | Noted
2025-05-17 20:59 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | Skull Tracker | https://drive.google.com/open?id=18RSoQuoJwxZBVT75gRIVRERxyt5DHmPd | 2025-05-13 | Rohan Ghosh | Dr. Bhubaneswara R & Dr. Venkatesh Yedulla | Cranial |  |  | 0.22 | 3.99 | 0.66 | https://drive.google.com/open?id=1jpxtYpuUlYsequ-slbdXukArSDSV3NLM | during the intraoperative period, while the patient was fully draped, an unexpected movement occurred where the patient slightly lifted his head and moved it from side to side. Simultaneously, a nursing staff member attempted to stabilize and reposition the patient’s head. This sequence of events raises the possibility that the skull tracker reference frame may have experienced minor displacement. | Dr. Bhubaneswara Raju provided mixed feedback regarding the use of the NeuroSaphireX system for this particular surgery. While the system initially demonstrated excellent accuracy and performance during setup and registration, the intraoperative disruption impacted the continuity of navigation reliability. He acknowledged the system’s strong performance during setup and the effectiveness of the Skull Tracker in a frameless scenario. However, he emphasized the need to ensure that anesthesia-related contingencies are managed to avoid accuracy loss in future cases. | https://drive.google.com/open?id=1ImuKpYEWq4MxfgoMbLeIojJclhkfl31g | No | It is important to note that the skull tracker was not manually repositioned during the procedure. However, during the intraoperative period, while the patient was fully draped, an unexpected movement occurred where the patient slightly lifted his head and moved it from side to side. Simultaneously, a nursing staff member attempted to stabilize and reposition the patient’s head. This sequence of events raises the possibility that the skull tracker reference frame may have experienced minor displacement. Another potential contributing factor could be suboptimal fixation of the skull tracker. It is suspected that the fixation screw may not have fully penetrated the cranial bone, which might have allowed the tracker block to shift slightly during patient movement. Both scenarios could have contributed to the observed decline in navigation accuracy. | during the intraoperative period, while the patient was fully draped, an unexpected movement occurred where the patient slightly lifted his head and moved it from side to side. Simultaneously, a nursing staff member attempted to stabilize and reposition the patient’s head. This sequence of events raises the possibility that the skull tracker reference frame may have experienced minor displacement. | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1S78AQQKKShEzhus894PxvLosi7am_4lv | https://drive.google.com/open?id=1YA2SQWGEJLg3dHfsr7rzVk9I78gGgh2L | [11 empty] | Noted
2025-05-17 21:33 | rohan@rxoom.com | Rohan Ghosh | Tx Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=1QqyupoTMRcWQv9jrQWk-r4ZRMk9frgcb | 2025-05-16 | Rohan Ghosh | Dr. Chandra Shekar Tavusetty | CT-Spine |  | 0 | 0.08 | 4.93 | 1.16 | https://drive.google.com/open?id=1Li3tAPe440unMGIfFTRCxpyK1FPU2EFE | During the spine navigation demo, the surgeon followed the standard registration protocol and uniformly marked six fiducial points on the T4 and T5 vertebrae. However, upon switching to navigation mode, the 3D model containing the marked fiducial points unexpectedly disappeared from the system interface. As a result, the surgeon was required to recall and re-register the fiducial points from memory, effectively performing the registration without visual guidance. Despite five repeated attempts, the system consistently displayed a surface accuracy error exceeding 4 mm, leading to the discontinuation of navigation support for the remainder of the procedure. | The surgeon has expressed interest in conducting additional cases using the navigation system before providing final feedback. He considers the current experience as part of his learning curve with the technology and prefers further hands-on exposure to fully evaluate the system’s performance. |  | No | During the spine navigation demo, the surgeon followed the standard registration protocol and uniformly marked six fiducial points on the T4 and T5 vertebrae. However, upon switching to navigation mode, the 3D model containing the marked fiducial points unexpectedly disappeared from the system interface. As a result, the surgeon was required to recall and re-register the fiducial points from memory, effectively performing the registration without visual guidance. Despite five repeated attempts, the system consistently displayed a surface accuracy error exceeding 4 mm, leading to the discontinuation of navigation support for the remainder of the procedure. | upon switching to navigation mode, the 3D model containing the marked fiducial points unexpectedly disappeared from the system interface. As a result, the surgeon was required to recall and re-register the fiducial points from memory, effectively performing the registration without visual guidance. Despite five repeated attempts, the system consistently displayed a surface accuracy error exceeding 4 mm, leading to the discontinuation of navigation support for the remainder of the procedure | NA | NA | NA | I kindly request the technical team to review this issue and explore a possible enhancement to the system—specifically, enabling the visibility of both the model and fiducial points even after entering navigation mode. This will support the surgeon in verifying fiducial accuracy during the registration process and help avoid similar challenges in the future. | The surgeon has expressed interest in conducting additional cases using the navigation system before providing final feedback. He considers the current experience as part of his learning curve with the technology and prefers further hands-on exposure to fully evaluate the system’s performance. | https://drive.google.com/open?id=1NhkUAT3GSUyH7aBPzqiCuOkiIiTCPabI | https://drive.google.com/open?id=1UYFFLjsXewODqdxyzBluWayNNXw4Tx0k | [11 empty] | Noted
2025-05-19 16:33 | gurjit758.232@gmail.com | Gurjit Singh | Paras Hospital | Gurgaon | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=13vponDsnvnFuqkhgjiAbhfoPq6y31yVA | 2025-05-19 | Gurjit Singh | Dr. Ravi | Cranial |  |  | 0.14 | 0.99 | 0.31 | https://drive.google.com/open?id=1pQHHDnV8DsEM7XXsqCaJmijyXyL49C3Z | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1oBV8pPhP_fcViDCdUW9rqLpWQ858PjGd | [12 empty] | Noted
2025-05-20 17:03 | janhavi@rxoom.com | Janhvi Barhate | P. D. Hinduja hospital | Mumbai | NS TRL 04 DB 101 | Hinduja system | 1.2.24 | NA | https://drive.google.com/open?id=1b6BuK7RiDbbUeEk2-Ur2I2Uv2NRD_upC | 2025-05-15 | Janhavi Barhate and Manoj Dhavale | Dr. Sameer Bhargav | Other | ENT |  | 0.24 | 5.05 | 0.55 | https://drive.google.com/open?id=1ifKeudqFsqPvZIUhiHxLx13fmOTFo0_r | NA | Bring something alternative to skull tracker |  | Yes |  |  | NA | Na | NA | Bring alternative to skull tracker | NA | https://drive.google.com/open?id=1jB611RhcVW98S5Ts1F2DgA3ZMu4ppeSc | [12 empty] | Noted
2025-05-21 11:16 | janhavi@rxoom.com | Faiz Khan | P D Hinduja hospital | Mumbai | NS TRL 01 MF 101 |  | 1.2.24 | Na | https://drive.google.com/open?id=1-J03XdTkS6jFOAUwSl-50IlZNo5csqSN | 2025-05-20 | Janhavi Barhate | Dr. b k mishra | Cranial |  |  | 0.24 | 5.25 | 0.42 | https://drive.google.com/open?id=1lQeLp8dVHdjX-uG4_9BXBniBerf53bml | Na | Na |  | Yes |  |  | Na | Na | Na | Na | Na | https://drive.google.com/open?id=19mL-Neetg-yJbYshqf-Bmu7XMoEpuM2M | [12 empty] | Noted
2025-05-21 11:20 | janhavi@rxoom.com | Faiz Khan | P. D. Hinduja hospital | Mumbai | NS TRL 01 MF 101 |  | 1.2.24 | Na | https://drive.google.com/open?id=1bwMc5WEzpAbvwgioRBUTvaqjd4kQxWy9 | 2025-05-20 | Janhavi Barhate | Dr. B. K. Mishra | Cranial |  |  | 0.15 | 8.39 | 0.53 | https://drive.google.com/open?id=1NPh6uI73gr1DLaCrVJca6WPzOLPpb0-N | Na | NA |  | Yes |  |  | Na | NA | Na | Na | Na | https://drive.google.com/open?id=1yoNr4_kwNZi5rMjlURp4MlalPh7m4aJB | [12 empty] | Noted
2025-05-22 10:10 | hemarshitha@rxoom.com | Hemarshitha H. S. | Aster MIMS Hospital | Calicut | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1E7oUS8JAYeypfSgrOh337cHg3DyaP8XI | 2025-05-16 | Hemarshitha HS | * Dr. Murali Krishna * Dr. Shreekumar | Cranial |  |  | 0.36 | 4.59 | 0.54 | https://drive.google.com/open?id=1pJuVHrb_-JXeNEk4bbNRY74D8W8WCkED | NA | The surgeon expressed full satisfaction with the navigation system’s performance, workflow, and accuracy. |  | Yes |  |  | Suggestion for Improvement Post-fiducial marking, the fiducial points should be made visible on the 3D model during registration to avoid confusion and enhance workflow clarity. | NA | NA | NA | NA | https://drive.google.com/open?id=1pewNvWDWGuifdr678zFVSBZIuq8ZVg5V | [12 empty] | Noted
2025-05-26 11:42 | trushalgulhane1547@gmail.com | Trushal Gulhane | Sahyandri hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | Stickers | https://drive.google.com/open?id=1EZiZHO03FU_svFr6Y6w3XWSX_KZaf1_R | 2025-05-22 | Trushal Gulhane | Dr. Sameer futane | Cranial |  |  | 0.4 | 3 | 1.12 | https://drive.google.com/open?id=1Ng2yECWdiekrQ_5CYvRFfTFysRXKwbjD | NA | Yes |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1aC9d7nKePJQclWqAy8eWLx5Mo2XBnP8t | [12 empty] | Noted
2025-05-26 11:52 | trushalgulhane1547@gmail.com | Trushal Gulhane | Sahyandri hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | Stickers | https://drive.google.com/open?id=1AZkBj0rVb1OEUUlj9wGvcOAe92uqLqzv | 2025-05-23 | Trushal Gulhane | Dr. Rahul modgi | Cranial |  |  | 0.2 | 4 | 0.4 | https://drive.google.com/open?id=1wGxhnz8EOArLxyLbcgJqs2fbP-vUm-xC | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1ABDTmEiBkS8LnSXn0xv8FhDzmSNfqQr7 | [12 empty] | Noted
2025-05-29 13:50 | gurjit758.232@gmail.com | Gurjit Singh | Paras Hospital | Gurgaon | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1Ds2-eDPeNVxshGPO-O7Ytds4DP1yB1TC | 2025-05-29 | Gurjit Singh | Dr. Raghav | Cranial |  |  | 0.28 | 4.35 | 0.59 | https://drive.google.com/open?id=1VXf5_ufAf6qer8Kowm9BckjH-SglBRpL | NA | Okay Accuracy |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=129EhfHe132sDy4yCA7ikivDsFP5oLSTi | [12 empty] | Noted
2025-05-29 14:09 | trushalgulhane1547@gmail.com | Trushal Gulhane | Sahyandri hospital | Nashik | NS TRL 04 DB 101 | Hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=1uef9ZpSzTz_xH8IWgV0_sMM98TiKIcHp | 2025-05-29 | Trushal Gulhane | Dr. Sameer futane and Dr rahul modgi | CT-Spine |  | 20 | 0.17 | 3 | 0.8 | https://drive.google.com/open?id=1Hr-599n5VQbWwMX8LdVd9vZs1JafNNBX | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=14akcBrTI5XZP0d4YOv3aQzXjFEwwu5lm | [12 empty] | Noted
2025-05-29 17:45 | hemarshitha1802@gmail.com | Hemarshitha H. S. | KMCT Medical College | CALICUT, Mukkama | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1q9MTW8jxg6YW1EMx_7Fum6ZGnKZSkOcL | 2025-05-28 | Hemarshitha H S | Dr Bimal Govind | Cranial |  |  | 0.18 | 6.94 | 1.16 | https://drive.google.com/open?id=1lM74ie4SxKWSSvAxLZAjFS6ta6ydP8fR | Preoperative MRI data was utilized for navigation purposes. However, multiple issues were identified with the dataset, including: * Incomplete capture of anatomical landmarks, notably the nose tip * Significant disturbances in the 3D dataset despite adherence to the prescribed acquisition protocol * The concerns were immediately brought to the attention of the surgeon, Dr. Bimal Govind, who then consulted with the radiologist. It was acknowledged that errors occurred during data acquisition. Despite the limitations, the surgeon opted to proceed with the available dataset. | Dr. Bimal Govind expressed complete satisfaction with the system. He noted that the registration process was smooth and effective, even with the patient in the prone position. The surgeon highlighted the system's performance, accuracy, and ease of use as major strengths. |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1MyxKI5xcHFU3qX5fgkfw8f5CQhNJulhx | [12 empty] | Noted
2025-05-29 18:57 | rohan@rxoom.com | Rohan Ghosh | Tx Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | V1.2.24 | NA | https://drive.google.com/open?id=179yaDO2QiWm4lKXbSF2VQNGRXeAg0On6 | 2025-05-25 | Rohan Ghosh | Dr. Naresh Kumar Gajjala & Dr Chandra Shekar Tavisetty | Cranial |  |  | 0.09 | 4.12 | 0.69 | https://drive.google.com/open?id=16WUkm8XD7TkNzZ6T_B3ZyiADKWgofJQQ | 1. During the procedure with Dr. Naresh, a horseshoe headrest was utilized instead of a 3-pin fixation system. I expressed concern that the lack of rigid head fixation might compromise the accuracy of the optical tracking navigation system. 2. Also prior to surgery, I suggested utilizing our skull tracker instrument to optimize accuracy. However, the surgeon opted not to use it, confident that accuracy would be maintained despite draping. The surgeon's confidence likely stemmed from the tumor's size and superficial location, where minor errors (1-2 mm) were deemed clinically insignificant. Fortunately, accuracy remained acceptable throughout the procedure. | Throughout the surgery, the NeuroSaphireX system performed reliably without any technical disruptions. Dr. Naresh Kumar Gajjala expressed his satisfaction with the system’s user interface, speed of patient data loading, and the efficiency of the registration process. He noted the system’s ease of use and smooth operation. Dr. Gajjala also expressed interest in utilizing the NeuroSaphireX system in upcoming biopsy cases, where accuracy is critical, to further evaluate the system’s performance in more complex scenarios. | https://drive.google.com/open?id=1QPdwSqKVVnOj4k3j-yPPcnnuto5_zhOS | Yes |  |  | NA | NA | NA | NA | Dr. Naresh Kumar Gajjala expressed his satisfaction with the system’s user interface, speed of patient data loading, and the efficiency of the registration process. He noted the system’s ease of use and smooth operation. Dr. Gajjala also expressed interest in utilizing the NeuroSaphireX system in upcoming biopsy cases, where accuracy is critical, to further evaluate the system’s performance in more complex scenarios. | https://drive.google.com/open?id=1hniVCeSNoNHEmcVN1GNyALsbq-FR6k6S, https://drive.google.com/open?id=1DiGlBpNBx21QJDVTm2ldfGRHKZxTDjDB | https://drive.google.com/open?id=1ZXSq0kwwAT14JJ7gMk98pVlOSwBJ9HXR | [11 empty] | Noted
2025-05-30 20:55 | janhavi@rxoom.com | Faiz Khan | Hinduja Hospital | Mumbai | NS TRL 01 MF 101 |  | V1.2.24 | NA | https://drive.google.com/open?id=1bGVSXLUu61KurRkh-aZBnfug9Z70BZ4G | 2025-05-30 | Janhavi Barhate | Dr. B. K .mishra | Cranial |  |  | 0.14 | 4.09 | 0.41 | https://drive.google.com/open?id=1OFLnF9ZjOhrNVrXepM9QsMnWzhJth8fP | Na | Na |  | Yes |  |  | Na | Na | Na | Na | Na | https://drive.google.com/open?id=1BP5xxR1pgGmODa1P9qaTZgbVjRk52T20 | [12 empty] | Noted
2025-05-31 19:42 | garadsuraj567@gmail.com | Faiz Khan | Max Neuro Hospital | Pune | NS TRL 02 MF 203 |  | R1.2 | NA | https://drive.google.com/open?id=1F_FyEAo0Q2D0xCPXPo-VKmrJxezyIw4j | 2025-05-31 | Suraj garad | Dr. Avishkar | Cranial |  |  | 0.22 | 4.11 | 0.98 | https://drive.google.com/open?id=1ieZsdgbuIxDnWPhqt2JeeXjHDnvp6ZzG | NA | NA |  | Yes |  |  |  | NA | NA | NA | NA | https://drive.google.com/open?id=1A7rBjzeDpmUsMtxFGltNR6bN_BYlSvyb
2025-05-31 21:37 | hemarshitha1802@gmail.com | Hemarshitha H. S. | KMTC Medical College | CALICUT, Mukkama | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1UQhD1dxvR856hBYmrPfXnbsL3K18kERN | 2025-05-30 | Hemarshitha HS | Dr. Bimal Govind | Cranial |  |  | 0.18 | 5.07 | 0.72 | https://drive.google.com/open?id=11wMcrBPcvlpEz82qLqq_IJoU9RdW8CvT | No | Dr. Bimal Govind expressed complete satisfaction with the performance of the navigation system. He noted that the system was intuitive and user-friendly. Additionally, he emphasized that the registration process was simple to execute even with the patient positioned prone. | https://drive.google.com/open?id=12Y7ZCwAUsY7k2MigUewgUkGkMizWofi2 | Yes |  |  | No | No | No | NA | NA | https://drive.google.com/open?id=116IhKfTsIGHxMMaZKcETINqs77OKmhFr
2025-06-04 16:44 | hemarshitha1802@gmail.com | Hemarshitha H. S. | KMCT Hospital | CALICUT, Mukkama | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1qIFlenXFiLXILocqGfakslupgDN6330D | 2025-06-03 | Hemarshitha HS | Dr. Bimal Govind | CT-Spine |  | 10 | 0.27 | 4.12 | 1.09 | https://drive.google.com/open?id=1gh_hbGjVp19-8JnXIevXXgEF0mlTpIio | During our recent use of the CT clamp, we observed an issue with its stability. The clamp was not holding steady and was continuously shifting from its position. This instability appears to be caused by the large size of the clamp, which creates excess leverage on the base and results in movement during use. | Dr. Bimal Govind expressed complete satisfaction with the system's performance, accuracy and confirmed that the technology contributed positively to the surgical workflow and outcome which significantly aided in the precision of screw placement and overall surgical navigation. |  | Yes |  |  | CT clamp is Big and because of Big size, The clamp was not holding steady and was continuously shifting from its position. This instability appears to be caused by the large size of the clamp, which creates excess leverage on the base and results in movement during use. | NA | System need to repaint. its looking so dull because of scratches. | NA | NA | https://drive.google.com/open?id=14aQreJ0up2PVaKeTAWzPV87ftOyrK3aj
2025-06-05 16:44 | dhavalepatil1100@gmail.com | Faiz Khan | Hinduja hospital | Mumbai | NS TRL 01 MF 101 | Hinduja hospital system | 1.1 | NA | https://drive.google.com/open?id=1uMJDAaKVMTD_0PSWRC_eVFumYHEfQfM0 | 2025-06-02 | Manoj dhavale | Ansul warade | Cranial |  |  | 12 | 3.5 | 0.55 | https://drive.google.com/open?id=1kPqc_IIsIBs5gIM9LNkBrnkrT9QEubVT | NA | Accuracy is good |  | Yes |  |  | Na | NA | NA | NA | NA | https://drive.google.com/open?id=1UFMaYxh1tyGFN6eJyjKahsELIuHa9OYK
2025-06-05 17:00 | dhavalepatil1100@gmail.com | Faiz Khan | Darsh hospital | Gulbarga | Other | Gulbarga system | 1.1 | NA | https://drive.google.com/open?id=1zoRdH5dCc4h4ZeTwZ5GfR8M2Toj884zK | 2025-05-25 | Manoj dhavale | Shashank Ramdurga | CT-Spine |  | 50 | 0.15 | 2.5 | 0.56 | https://drive.google.com/open?id=1U8ApJzey27tb9yrHs1-Cs2U2FJjks5YS | NA | Accuracy is good | https://drive.google.com/open?id=1yBKn7ZLyN_WgJf601O47q4uotyhxROCJ | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1FVp4Etj8B6kOiHXEwOYmNxsBaEEbot3Y
2025-06-16 11:39 | hemarshitha1802@gmail.com | Hemarshitha H. S. | KIMS Hospital, Kannur | Kannur, Kerala | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1nr4XYopAINrKb2tQ7kyfkv17H7VaMD6l | 2025-06-14 | Hemarshitha HS | Dr. Mahesh Bhat | Cranial |  |  | 0.08 | 7.79 | 0.71 | https://drive.google.com/open?id=1qN8zN0Fz7VRO2akPNLzofEQ4Db-JanK3 | NO | Dr. Mahesh Bhat expressed full satisfaction with the performance of the navigation system. He found the system to be intuitive and user-friendly. Despite the patient being positioned prone, the registration process was straightforward and executed without any difficulty | https://drive.google.com/open?id=1zuVzwADt8w7-g6iM0ArdVObO5RZlxR2K | Yes |  |  | NA | NA | System need to repaint. its looking so dull because of scratches. | NA | NA | https://drive.google.com/open?id=1e9PoSG6huqOucm9d0tvH2igm2nZDwBpQ
2025-06-16 22:15 | gurjit758.232@gmail.com | Gurjit Singh | Wellsun Medicity | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1wrGJOimOUnMas8WMxQd-b3RBevdwLaTA | 2025-06-16 | Gurjit Singh | Dr. Bhawan Nangarwal | Cranial |  |  | 0.18 | 3.33 | 1.03 | https://drive.google.com/open?id=1UPLFIgR2dzqZFSzeVKXIcViSbYFUrY74 | NA | Good accuracy |  | Yes |  |  | NA | NA | Difficult to control while moving trolley in floor/OT. It Don’t easily move the way we want. Makes it difficult to carefully handle the system while moving/shifting the whole system in and out of OT while system take a lot of push power and go in unwanted direction. | A drawer in system to keep general tools easily accessible and to provide a better place to keep some number of tools while doing demo. | NA | https://drive.google.com/open?id=1Y7-5f_J5OD-GNsR00StkzIeYuNSL4jg3
2025-06-20 15:44 | garadsuraj567@gmail.com | Faiz Khan | Noble hospital | Pune | NS TRL 02 MF 204 |  | 1.2.24 | Mayfield adaptor | https://drive.google.com/open?id=1_dnDr0SA4ppHW1VzJVCeVTShg6Lbn-V4 | 2025-06-19 | Suraj garad | Dr. R Chandrashekhar | Biopsy, Any (Cranial, Full body) |  |  | 0.1 | 3.04 | 0.34 | https://drive.google.com/open?id=1gK7qgxhzk1__9HceLmjL9wacC-XR_ulv | Na | Na |  | Yes |  |  |  | NA | NA | NA | NA | https://drive.google.com/open?id=1-BiUmo9MsBMmFsS2cRDw3TxEzwx0lwxO
2025-06-20 20:20 | gurjit758.232@gmail.com | Gurjit Singh | Max Hospital | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1HB49STYuO44EGcWo-bZnZHOi9BwWJ6w1 | 2025-06-20 | Gurjit Singh | Dr. Mazhar Husain and Dr Siddhart | Cranial |  |  | 0.24 | 3.76 | 0.46 | https://drive.google.com/open?id=17I5TgL2vh0yAjSqhVVWCRgP5hRYaoS-8 | Not able to fix skull tracker with mini plate due to hospital did not have bigger skrew and the screw driver they had didn’t matched with our skrew hread. | Dr Mazhar Husain Good Navigation experience and Assistant Dr. Siddhart |  | Yes |  |  | NA | NA | Trolly moving control issue on floor. | NA | Keep the matching screwdriver of screws too. | https://drive.google.com/open?id=1SjvANmMHHYxXWWuVA_Hh2mzEGYy3bjCd
2025-06-21 13:32 | trushalgulhane1547@gmail.com | Faiz Khan | Hinduja hospital | Mumbai | NS TRL 01 MF 101 | Hinduja Hospital | V1.2.24 | NA | https://drive.google.com/open?id=1aUTQPywaf8QHIh3_DvX0NMv-yL5YRNpO | 2025-06-20 | Trushal Gulhane | Dr B. K. Mishra | Cranial |  |  | 0.14 | 3 | 0.5 | https://drive.google.com/open?id=1_ivRz5tMNnzobfYS1PsVtfnApfsTV3kG | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=11J-b_UQiu7GdfqyX0MlOhmBFRRzFsxJN
2025-06-21 15:52 | dhavalepatil1100@gmail.com | Faiz Khan | Noble hospital | Pune | Other | Noble Hospital | 1.2 | NA | https://drive.google.com/open?id=1C6mQhhRYtLFK1e-BO5f7FP3Iy1ehhXrN | 2025-06-20 | Manoj dhavale and aadarsh | Chandrashekhar raman | Cranial |  |  | 0.1 | 2.3 | 1.2 | https://drive.google.com/open?id=1uTVgT2Ll5gLG97OBm-AkSewH_fAnzn3b | Na | NA |  | Yes |  |  |  | Na | Na | Na | Na | https://drive.google.com/open?id=1Cil4Kp814BirY6BYxm3chako947MC8Xl
2025-06-23 01:18 | rohan@rxoom.com | Rohan Ghosh | Sentini City Hospital | Vijaywada | NS TRL 03 SE 101 |  | 2.24 | Lalya arm and adapter | https://drive.google.com/open?id=178g3afSrqvOqbcMf6s3gBlNV0b1B1qgR | 2025-06-17 | Rohan ghosh | Dr. MD Ameen | Cranial |  |  | 0.2 | 3.93 | 1.2 | https://drive.google.com/open?id=1ZtpRY7XDUCegZsp2zzNsyhZ8qoC9Jwkt | NA | case involved verification of our frameless navigation system's accuracy. The region of interest (ROI) was located just beneath the skull, and the tumor was superficially positioned. Dr. Ameen performed the registration process himself and expressed satisfaction with the system's accuracy, confirming its reliability for clinical use. This test case demonstrated the system's precision in a real-world setting. | https://drive.google.com/open?id=1dQkMaUhadfGcrcaZ1JBxFggo49tCIQt7 | Yes |  |  | NA | NA | NA | NA | case involved verification of our frameless navigation system's accuracy. The region of interest (ROI) was located just beneath the skull, and the tumor was superficially positioned. Dr. Ameen performed the registration process himself and expressed satisfaction with the system's accuracy, confirming its reliability for clinical use. This test case demonstrated the system's precision in a real-world setting. | https://drive.google.com/open?id=1_k-BF1y3vzkBK0Hsn59klx-xYLULgIX8 | https://drive.google.com/open?id=19lfX0SkwjQj4ILBlMGZ36Pah9cx67sN4
2025-06-24 22:44 | gurjit758.232@gmail.com | Gurjit Singh | Max Hospital | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1A3tZI-8dTi9lOphrkXCKiyFxcxK4X5me | 2025-06-24 | Gurjit Singh | Dr. Manu Rastogi | Cranial |  |  | 0.14 | 3.5 | 0.36 | https://drive.google.com/open?id=1_AyOf25uebt-sBzCClS-BOHmrxrJFDX- | NA | Like to check accuracy in more cases. |  | Yes |  |  | MRI Data received, but the dataset were like 2 slices, 3 slices but up one loading the full model is there. Count number of slices in dataset library didn’t matched with slices showing in views after loading dataset. | NA | NA | NA | NA | https://drive.google.com/open?id=1Jh25De9ZH9QBW5_k6Ix1fmnqH7Roed-h, https://drive.google.com/open?id=1SAlmfi_GwPAG9B54a4PwWPJVdK557rN6
2025-06-26 10:33 | trushalgulhane1547@gmail.com | Faiz Khan | Hinduja hospital | Mumbai | NS TRL 01 MF 101 | Hinduja hospital | V1.2.24 | NA | https://drive.google.com/open?id=1J_ORJeF_xefWisI5Y_t45hSR7m36ic86 | 2025-06-24 | Trushal Gulhane | B k mishra | Cranial |  |  | 0.1 | 3 | 0.5 | https://drive.google.com/open?id=1gEojWi9EFelNZn4NdIghOfIncZhxwBGr | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1NBrRNJ71-NG659IMm5wg_PdHUiZUT1ku
2025-07-04 13:54 | trushalgulhane1547@gmail.com | Faiz Khan | Hinduja hospital | Mumbai | NS TRL 01 MF 101 | Hinduja hospital system | V1.2.24 | NA | https://drive.google.com/open?id=1xi3zv1Fgf8IgnMK0-6C1NrHpLRx__9zc | 2025-07-03 | Trushal Gulhane | Sameer Bhargav | Other | ENT |  | 0.1 | 1.29 | 1.44 | https://drive.google.com/open?id=1WiG-5g5RzYnmAOIPU1svP-yrYUpGe6dO | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=12Fq1Rq7AOKnTnuk06otCMEZc70P4I3jZ
2025-07-06 10:04 | trushalgulhane1547@gmail.com | Trushal Gulhane | Sahyandri hospital | Nashik | Other | Middle hexagonal middle pipe | V1.2.24 | NA | https://drive.google.com/open?id=1-ntMq8Ew0tlcCJOK72l0H-DVdtKOSPZu | 2025-06-04 | Trushal Gulhane | Dr Sameer futane | Cranial |  |  | 0.2 | 3.4 | 0.9 | https://drive.google.com/open?id=1GS6VZxg3MDsgu4lwIUw_2RcM-O-KJ7BI | NA | NA |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1HGTOZTwuABG-YtyHBJThQbaD23ZsF2kW
2025-07-07 21:17 | gurjit@rxoom.com | Gurjit Singh | Max Hospital | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=170qyXgMpBTKglzjyX3vZmumLKmG0n_pU | 2025-07-07 | Gurjit Singh | Dr. Manu Rastogi and Dr. Sirddhart | Cranial |  |  | 0.42 | 6.78 | 0.79 | https://drive.google.com/open?id=1iljRrDseRLxup6uK2Ua1TxMnYu1mqoI8 | Fiducial accuracy was coming high again and again. | Good accuracy and concept like Medtronic |  | Yes |  |  | In 3D MRI patient face seems puffed than actual | NA | NA | NA | NA | https://drive.google.com/open?id=1yNR1nA5UxKtIGHEqt9Sf5YL9w31jlBfS
2025-07-08 01:19 | rohan@rxoom.com | Rohan Ghosh | AIIMS, Mangalagiri, AAP | Mangalagir, Amravati, AP | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=194z9jcjiV-zkeb7-i-MGgbwiuRnDROZG | 2025-07-07 | Rohan Ghosh | Dr. Rajasekhar R | Cranial |  |  | 0.1 | 4.18 | 0.59 | https://drive.google.com/open?id=1FvIna8Rm0-dRJJax3Hy3cQSd04hNdnJ0 | NA | 1. Dr. Rajshekar emphasized that integrating the navigation system with surgical microscopes and operative instruments such as suction devices would be a significant advancement. He noted that many competing systems already offer such integrated features, which serve as key differentiators and unique selling points in their marketing strategies. Incorporating similar capabilities into our system could enhance surgical efficiency and broaden clinical applicability, making the NeuroSaphireX system more competitive and attractive to end users. 2. Dr. Rajshekar highlighted that in complex cases involving far lateral extension and deep-seated tumors, other navigation systems typically exhibit an error margin of around 2mm when the hand probe is advanced near the tumor. In contrast, the NeuroSaphireX system demonstrated significantly higher precision, maintaining an error margin of less than 0.50mm. He emphasized that this level of accuracy is clinically significant and enhances surgical confidence during critical tumor resections, particularly in anatomically challenging areas. | https://drive.google.com/open?id=122ngPSoW8Wjw0PYYO01Fgwj4xiUIsZP0 | Yes |  |  | NA | NA | NA | My input is , that, Dr. Rajshekar emphasized that integrating the navigation system with surgical microscopes and operative instruments such as suction devices would be a significant advancement. He noted that many competing systems already offer such integrated features, which serve as key differentiators and unique selling points in their marketing strategies. Incorporating similar capabilities into our system could enhance surgical efficiency and broaden clinical applicability, making the NeuroSaphireX system more competitive and attractive to end users. | 1. Dr. Rajshekar emphasized that integrating the navigation system with surgical microscopes and operative instruments such as suction devices would be a significant advancement. He noted that many competing systems already offer such integrated features, which serve as key differentiators and unique selling points in their marketing strategies. Incorporating similar capabilities into our system could enhance surgical efficiency and broaden clinical applicability, making the NeuroSaphireX system more competitive and attractive to end users. 2. Dr. Rajshekar highlighted that in complex cases involving far lateral extension and deep-seated tumors, other navigation systems typically exhibit an error margin of around 2mm when the hand probe is advanced near the tumor. In contrast, the NeuroSaphireX system demonstrated significantly higher precision, maintaining an error margin of less than 0.50mm. He emphasized that this level of accuracy is clinically significant and enhances surgical confidence during critical tumor resections, particularly in anatomically challenging areas. | https://drive.google.com/open?id=1HHoFj3fuVegcbK6ZcOna1ghKBN4rx6Ld, https://drive.google.com/open?id=1Y9kqy-cqy1EjFPCuttfHIIqlNLvBfgrD, https://drive.google.com/open?id=1XBPnBEr4PGvgvbV3pyrarNIa1xJ9MtNM, https://drive.google.com/open?id=17uHluxK3zfG86HUNTg9h1klFa2HK918J | https://drive.google.com/open?id=15Uh0OaoIbj0v4s6pV3zwtOfxMDO3O1bR
2025-07-08 13:58 | dhavalepatil1100@gmail.com | Mumbai Distributor | Hinduja hospital | Mumbai | NS TRL 01 MF 101 | Hinduja hospital | 1.2.24 | NA | https://drive.google.com/open?id=1Hk29fJqo5YQUj9IV0t4ZdqaKt5rVsCmO | 2025-07-07 | Manoj Bhausaheb dhavale and gajanan | Jayshankar narayan | Cranial |  |  | 0.16 | 6 | 1.11 | https://drive.google.com/open?id=11CxYtKbaREFsfxYTzq2bFSJqJUPHVwMa | Pateint data is not capture properly ,need to do multiple re-registration | Na |  | Yes |  |  | Na | Na | Na | Na | Na | https://drive.google.com/open?id=1RUUxoiev5GNuaLDB_N9h5_wNiPgFcTfR
2025-07-09 21:40 | gurjit758.232@gmail.com | Gurjit Singh | Wellsun Medicity | Lucknow | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1zihE9v_b0AkLuUJsF9iKcgUGWZGM0sfT | 2025-07-09 | Gurjit Singh | Dr. Bhawan Nangarwal | CT-Spine |  | 30 | 0.21 | 1.81 | 0.92 | https://drive.google.com/open?id=1ITYWo0KVSTYE2p7qtzIqjVcvasGxTCGc, https://drive.google.com/open?id=1hZEQIbVcqU4mqucSOk_C_mi-kWBbkZfT, https://drive.google.com/open?id=15uWF36Jo-pb4TwKxoNTW1GHLVu6pfLdL |  |  |  | No | What suppose to happen:- After spine single level vertebrate registration. Surgeon supposed to check and confirm accuracy on that single vertebrae with help of C Arm or comparing visual anatomical landmark with navigation accuracy and then use the navigation guidance to perform pedicle screw fixcation and which help reduce C arm radiation exposure. What happened:- After spine single level vertebrate registration. Dr. tried to check on other spine level vertebra on keep informing that high accuracy will be on the single register vertebrae then he checks the top front spine and top back spinous process of register vertebrae landmark was accurate then he put probe on pedicle try confirm with visual and said it is not accurate I asked sir to confirm with C arm Dr started confirming level with c arm then asked me to provide pedicle finder. Pedicle finder was not available then Surgeon rejected use of navigation without navigable spine instruments and removed the spine clamp. | Not having multiple navigable spine instruments, CT navigation with spine calibration system with C-arm in our system required by Surgeon to perform spine surgery. | NA | NA | NA | Required a study spine reference clamp fixation, to attach reference firmly on spine. Design should be concise and easily adaptable while fixing on spine so that it obstruct less in surgery operating area. | Provide multiple navigable spine instruments, develop CT navigation with spine C-arm calibration system to match or reflect changes in pre operative CT spine data using intra operative X-ray to better perform spine surgery. | https://drive.google.com/open?id=1nCfmIxZJQgnArHRIiyyOu1s6s6yCTFDC |  |  | NA | Surgeon required pedicle finder and other instruments which help provide guidance while inserting pedicle screws | NA | NA | Surgeon said accuracy is good on pedicle | NA | NA | Surgeon required C-Arm integration with navigation to calibrate the spine or update the change in preoperative CT scan just like HRS navigation provide. Surgeon required pedicle finder and other instruments which help provide guidance while inserting pedicle screws just like HRS navigation and Medtronic provide in spine navigation.
2025-07-11 22:44 | rohan@rxoom.com | Rohan Ghosh | AIIMS Mangalagiri, Amravati, AP | Mangalagiri, Amravati | NS TRL 03 SE 101 |  | 2.24 | Lelya Aram & Adapter | https://drive.google.com/open?id=1pqrqYINAQD5XuLqiMs7bXQ-fLtQnrXIC | 2025-07-11 | Rohan Ghosh | Dr. Rajasekhar R | Cranial |  |  | 0.07 | 4.01 | 0.67 | https://drive.google.com/open?id=1jYHbj0rnD5km4Yge0GNe7NgBqIAUAMIC, https://drive.google.com/open?id=1JdQmV9-B4MUT0tpEXWELQJihzAJkMBXF, https://drive.google.com/open?id=1IQ3e9JAyXYdlwqLWgAfPZVs_RFDvuTF7, https://drive.google.com/open?id=17LpE_FrLwy9xSOubwJnEIVAY_u18jgPN, https://drive.google.com/open?id=1giEI9ImgAzpFnPiSixFf7mcvCSOm88Rq |  |  | https://drive.google.com/open?id=1b8D56ndBmTNem_WomBYCk3vit1QX-RR6 | Yes |  |  | Dr. Rajshekar emphasized that integrating the navigation system with surgical microscopes and operative instruments such as suction devices would be a significant advancement. He noted that many competing systems already offer such integrated features, which serve as key differentiators and unique selling points in their marketing strategies. Incorporating similar capabilities into our system could enhance surgical efficiency and broaden clinical applicability, making the NeuroSaphireX system more competitive and attractive to end users. | NA | NA | Based on Dr. Rajasekhar R’s feedback, I suggest adding instrument integration (e.g., microscope, suction tools) into the navigation system. This will match competitor features and improve clinical utility. Additionally, our current compact design and user-friendly interface should be retained during future design, logistics, and assembly enhancements. | Dr. Rajasekhar R appreciated the NeuroSaphireX system for its user-friendly interface, compact design, and easy-to-use workflow, including patient data loading and registration. He emphasized the need for future integration of surgical instruments into the system, suggesting it would significantly enhance functionality and align with global standards. | https://drive.google.com/open?id=13NydzipLMavfoAa6NVEThNmE7_022-Jf, https://drive.google.com/open?id=1WxxkvRjcEu-IZEzM0sO78O0inatt7xh8, https://drive.google.com/open?id=1C4OwzHp2SNlgwy-PuqN8DhpxfRmMv6CS, https://drive.google.com/open?id=1wCy3lcFD-lkw-vE2zab1x5ad1bvBNSI0 | https://drive.google.com/open?id=11FM8mNbriOiRSCGI5c1gYFU_tpBprmXj |  | NA | NA | NA | Doctor asked us to develop the integration of different instruments. | Strongly Agree | Strongly Agree | Strongly Agree | User-friendly, Compact device, Stunning display
2025-07-18 19:16 | gajananbitode606@gmail.com | Gajanan Bitode | Pd Hinduja hospital mumbai | Mumbai | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1DeIkIkUMQB3nXMzfiyMDAk6W7k0jcPHm | 2025-07-18 | Gajanan bitode | DR BK MISHRA | Cranial |  |  | 0.09 | 5.65 | 0.61 | https://drive.google.com/open?id=1L3hJ5e4y6YDkOydH3Bsq0SBKAw9mu7qc |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=19X_jz3ywE9wHS3jiej31-GwMG3M4DA4K |  |  | NA | NA | NA | Na | NA | NA | Na | Na
2025-07-21 23:25 | hemarshitha@rxoom.com | Hemarshitha H. S. | Daya General Hospital | Thrissur | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=10dMQU21PrlRTLU8Ian-R5Q8ukrVfhnDe | 2025-07-19 | Hemarshitha H S | Dr Balu Mohan | Cranial |  |  | 0.18 | 4.1 | 0.58 | https://drive.google.com/open?id=1maHosajg4Qi80OauQlpUiITmoG-LOwB- |  |  | https://drive.google.com/open?id=1lQEq6KvYvYP6-8Ws6UdhPR9NmkqV0JUH | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=15ZjIniBfc09RvF0H9bAeF4TooUkCB9UE |  |  | NA | NA | NA | NA | Good Accuracy throughout the Surgery | NA | Easy to use | Will need more demos
2025-07-21 23:31 | hemarshitha@rxoom.com | Hemarshitha H. S. | Brindawan Hospital | Mysore | NS TRL 02 MF 204 |  | V1.2.24 | NA | https://drive.google.com/open?id=1kHkDQDjElBOwwWYBY-paPzYcbgG7WMFw | 2025-07-21 | Hemarshitha HS & Krishna Kumar | Dr. Anil Sangli | Cranial |  |  | 0.18 | 3.98 | 1.76 | https://drive.google.com/open?id=12hqjrSHR4ckOVK59ZnTNu3ACWMzycIpJ |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1ggKbUOJvUgwrVmy6n3VRlgzGFvFyw9hR |  |  | NA | NA | NA | NA | Accuracy was good | NA | Easy to use | Need more demos on deep seated tumor & Skull base surgeries.
2025-07-22 22:39 | hemarshitha@rxoom.com | Hemarshitha H. S. | Brindavan hospital | Mysuru | NS TRL 02 MF 204 |  | 1.2.24 | NA | https://drive.google.com/open?id=1tWnKj3ApSsv0oT1WcYqsPcL3Z4wWO2Tf | 2025-07-22 | Hemarshitha H S & Krishna Kumar | Dr Sree Harsha | X-Ray Spine |  | 40 | 0.15 | 00 | 0.8 | https://drive.google.com/open?id=16hCVWhX3b8hGWaSvcsiLjK1Pg4wmlmIK |  |  |  | No | Supposed to accurate in both Lateral and AP images but it was accurate in AP not in lateral. | Might be data related issue. |  | NA | NA | NA | Provide a reference that can be fixed on Jamshidi needle | https://drive.google.com/open?id=15x8DV-vnf5B_7DQZda8_0kEdYtyd3N_X |  |  | NA | NA | The X-ray spine case was unsuccessful because, although all eight points were visible after image acquisition, the probe displayed accurate positioning in the AP view but appeared to be in the air in the lateral view. The issue persisted even after three attempts. | NA | In AP view probe was showing accurate but in lateral the it was showing in the air | NA | Easy to register | In percutaneous screw fixation cases, it would be helpful to provide a reference frame that can be fixed onto the Jamshidi needle, as it is otherwise difficult to use both the probe and the needle simultaneously.
2025-07-23 09:32 | gajananbitode606@gmail.com | Gajanan Bitode | P.d hinduja hospital mumbai | Mumbai | NS TRL 01 MF 101 |  | 1.2.24 | NA | https://drive.google.com/open?id=1P8_jl0zjj3D-r4-FhcQvv_xyVpSFS429 | 2025-07-22 | Gajanan | Dr b.k mishra | Cranial |  |  | 0.09 | 3.88 | 0.61 | https://drive.google.com/open?id=1YV1FtbTiO3ZxHxk4ocPqGrT3u902cYIS |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1kZLMZdGQYnpVsu_e1ErkC00Hiq6o1Ps2, https://drive.google.com/open?id=12P0brkwVWTI6RGIYrrq84Dwckg6opW9s |  |  | Na | Skull clam adapter was affected by corrosion and doctor also complaint regarding this | NA | NA | Na | Na | Na | Na
2025-07-29 01:14 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=1mO6NpaCb2CAatNl1s_lnCSjp-0n61wR3 | 2025-07-24 | Rohan Ghosh | Dr. Manik Prabhu, Dr. T.N Rao | Cranial |  |  | 0.07 | 3.52 | 0.59 | https://drive.google.com/open?id=1ZLxtLfPHlk_ELv4nlKqrgNeMx1_dH1EH |  |  | https://drive.google.com/open?id=1e5-zFH4SXddLUYCm5_Ap0eBo9irH3anm | Yes |  |  | NA | NA | NA | Doctor asked to develop EM Navigation system for pediatric neuro cases. | Doctor asked to develop EM Navigation system for pediatric neuro cases. | https://drive.google.com/open?id=17VD34onaRoCF5KHv8NeaHKLQE-Vqq3Uo | https://drive.google.com/open?id=1XFAa-Xj-IebMJ33AN61A9YyUl81523aO |  | NA | NA | NA | NA | Strongly Agree | Strongly Agree | Strongly Agree | Very easy to use
2025-07-29 01:32 | rohan@rxoom.com | Rohan Ghosh | Care Hospital, Banjara Hills | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | 6mm Screw, which use in cranioplasty case | https://drive.google.com/open?id=1cnKWrJKy-P-rs4Zld1UEqYudj6afs14p | 2025-07-28 | Rohan Ghosh | Dr. Ajit, Dr. Manik Pabhu, Dr. TN Rao | Cranial |  |  | 0.14 | 3.15 | 0.67 | https://drive.google.com/open?id=1f8MErnCGy2SYSgivOHY5LPyu13cguDFv |  |  | https://drive.google.com/open?id=17KtD3MzUgC-cOQ4lJUSgIJs1b4luu_Go | Yes |  |  | NA | NA | NA | NA | Doctor asked to develop EM Navigation system for pediatric neuro cases. | https://drive.google.com/open?id=1qYiNuJF7zvl3V-9csIcJI3K2OFKqhKih, https://drive.google.com/open?id=1FkbX8QuGCUrZ_80icn_De3ft9XSdh9EW | https://drive.google.com/open?id=1rh7ugX0vTgm7ED1mDr8ZO8Ewfl77ZsZB |  | NA | NA | NA | At the end of the surgery, a shift in the brain’s midline was observed, likely due to cerebrospinal fluid (CSF) loss or tissue relaxation. As a result, the navigation system displayed a deviation of approximately **2 mm**, particularly when the probe was placed near bony landmarks, indicating a loss of alignment due to brain shift. | Strongly agree | Strongly agree | Strongly agree | At the end of the surgery, a shift in the brain’s midline was observed, likely due to cerebrospinal fluid (CSF) loss or tissue relaxation. As a result, the navigation system displayed a deviation of approximately **2 mm**, particularly when the probe was placed near bony landmarks, indicating a loss of alignment due to brain shift.
2025-07-30 00:35 | hemarshitha@rxoom.com | Hemarshitha H. S. | Brindavan hospital | Mysuru | NS TRL 02 MF 204 |  | 1.2.24 | NA | https://drive.google.com/open?id=1YvKm--uFpaS5CxfUA5GCXs2Wrixd8VhR | 2025-07-29 | Hemarshitha H S & Krishna Kumar | Dr Anil Sangli | Cranial |  |  | 0.32 | 5.35 | 1.32 | https://drive.google.com/open?id=1pyALQH6K7eVaQrGybWGnKEIHlDgptxe6 |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1R-OSspOiteMRnuqGsiPcb41CLnzJHTe9, https://drive.google.com/open?id=1ZHZ93hqh8MR3EWwK74yFMjuBNik1ofbc |  |  | NA | NA | NA | NA | Accuracy is maintained consistently throughout the surgery | Good | Easy to merge CT and MRI dataset | Surgeon appreciated the skull tracker as a good option for pinless surgery
2025-07-30 12:14 | dhavalepatil1100@gmail.com | Manoj Dhawale | Bombay hospital | Mumbai | Other | EM system | 1.3.1 | No | https://drive.google.com/open?id=1xWZ0zTQQd-Q3M-vlHw4Qp-enLMt0JRhN | 2025-07-30 | Manoj Bhausaheb dhavale and gajanan bitode | Dr. Devpujari and Nimit shah | Cranial |  |  | 0.1 | 5 | 0.72 | https://drive.google.com/open?id=1qWIA9mt64LtoRAuMimUGxgtng3798Gnv |  |  |  | Yes |  |  | No | No | No | No | No | https://drive.google.com/open?id=1jmt5wA0qqUArNQGFAafpi5w9AFS6g0v9 |  |  | We are trying murge data it gate crash | Surgeon want sterile probe tip should be long | No | Probe and reference wire are coming in between to surgeon and OT staff, they want long probe | Accuracy was good | They want probe tip long , and want long wire | No feedback | NO
2025-08-01 17:45 | gajananbitode606@gmail.com | Gajanan Bitode | Bombay hospital | Mumbai | Other | EM system | 1 | NA | https://drive.google.com/open?id=18gJe-2hOC3b_OnKUtZaBwjjDapYDLjdF, https://drive.google.com/open?id=1tEDlxzgiuO-ixUQRpgbmgOptrLumusi2, https://drive.google.com/open?id=1NgUV-eLpOArXnUJrINK2iW6ukhrUMRxa | 2025-08-01 | Gajanan bitode, trushal gulane | Dr deopujari sir | Cranial |  |  | 0.09 | 4.73 | 0.71 | https://drive.google.com/open?id=15GjVYrykOZFcuHkgFCHJXiLjjr1DUZcR, https://drive.google.com/open?id=1PwnA4fTDo55oxXvnHjRVQzXIMbP2uOkk |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1igjp2mQhDWEWHrEx5CU_0TdPNtkZirBa |  |  | NA | NA | NA | NA | surgeons said error of 5/6mm | NA | NA | NA | NA
2025-08-04 21:48 | rohan@rxoom.com | Rohan Ghosh | Medicover Hospital, Hitech city | Hyderabad | NS TRL 03 SE 101 |  | 2.24 | NA | https://drive.google.com/open?id=1ZCt4CnM3y5m13RWg8_laUP_yeH9v5qSS | 2025-08-04 | Rohan Ghosh | Dr. Srikant Reddy R | Cranial |  |  | 0.08 | 3.68 | 0.7 | https://drive.google.com/open?id=1n2dtPi7Ql0GzCFpxDB_U8xzOtEaIFYgX |  |  | https://drive.google.com/open?id=1rx4o7EbBw2UKAe4WynV916mUQ_ubLh9t | No | During the craniotomy procedure under the supervision of Dr. Srikant Reddy R, the NeuroSaphireX system was used for intraoperative navigation. Registration was completed successfully prior to draping. However, upon utilizing the sterile navigation probe, the surgeon noted that the Coronal, Sagittal, and Axial planes were not centered on the display, which made it difficult to interpret all three views simultaneously with clarity. Midway through the surgery, the surgeon reported cerebrospinal fluid (CSF) leakage and a resulting brain shift, which compromised navigation accuracy. As a result, the system showed an alignment deviation of up to 5 mm, limiting its effectiveness in real-time tracking beyond this point. | cerebrospinal fluid (CSF) leakage and a resulting brain shift, which compromised navigation accuracy. As a result, the system showed an alignment deviation of up to 5 mm, limiting its effectiveness in real-time tracking beyond this point. | NA | NA | NA | NA | It was suggested to enhance the display size of CT/MRI images and enable synchronized movement of MRI slices with the navigation probe for better visualization. | https://drive.google.com/open?id=1r1wJ6iPqTjicopvrNOfgbXVjGOhdOoeA, https://drive.google.com/open?id=1zXcM5JUTMBPH7MbaTQSUGpNWTFnea4ca | https://drive.google.com/open?id=1lpBrvBICkVhhZHTwF-s_clJi6X3cp5ba |  | During the craniotomy procedure under the supervision of Dr. Srikant Reddy R, the NeuroSaphireX system was used for intraoperative navigation. Registration was completed successfully prior to draping. However, upon utilizing the sterile navigation probe, the surgeon noted that the Coronal, Sagittal, and Axial planes were not centered on the display, which made it difficult to interpret all three views simultaneously with clarity. | NA | Only CT Brain Thin data was provided for the case. Due to the absence of MRI or contrast-enhanced imaging, tumor visibility on the navigation screen was limited, making it challenging to clearly differentiate tumor margins during surgical planning and navigation. | NA | Didn't Update because CSF leakage & brain shift | Didn't Update because CSF leakage & brain shift | Didn't Update because CSF leakage & brain shift | Improve in size of CT/MRI films that display & MRI films should be move along with the probe | Surgery Type: Craniotomy Surgeon: Dr. Srikant Reddy R Navigation Registration: Completed by application specialist During the craniotomy procedure under the supervision of Dr. Srikant Reddy R, the NeuroSaphireX system was used for intraoperative navigation. Registration was completed successfully prior to draping. However, upon utilizing the sterile navigation probe, the surgeon noted that the Coronal, Sagittal, and Axial planes were not centered on the display, which made it difficult to interpret all three views simultaneously with clarity. Midway through the surgery, the surgeon reported cerebrospinal fluid (CSF) leakage and a resulting brain shift, which compromised navigation accuracy. As a result, the system showed an alignment deviation of up to 5 mm, limiting its effectiveness in real-time tracking beyond this point. The feedback highlights the importance of improved display calibration for multi-planar visualization and the potential need for adaptive brain-shift compensation features to maintain navigational precision in such intraoperative conditions. | Optical
2025-08-18 23:48 | hemarshitha@rxoom.com | Hemarshitha H. S. | Lakeview hospital | Belgaum | NS TRL 02 MF 204 |  | 1.2.24 | NA | https://drive.google.com/open?id=11s1da1XUsSVEMY-qCjECf6V42Bzn0mJe | 2025-08-18 | Hemarshitha H S | Dr Raviraj Gorpade | Cranial |  |  | 0.16 | 3.39 | 0.52 | https://drive.google.com/open?id=1_Fad-MyUf9yhkBqRe36-lZe3vDloZ8hf |  |  |  | Yes |  |  | Face One issue regarding the Skull Clamo Arm, Arm was not fixing to Three Pin Mayfiled because mayfiled gear was small, and even our Knob was not going inside the Gear because leanght was big of tightening Knob. Doctor also informed me to change the size if Arm and Knob. | NA | NA | NA | NA | https://drive.google.com/open?id=1m1uS-rXs5FQBgiHCl1fcxt3KNHdYNoMQ | https://drive.google.com/open?id=1-CqCFFh2ClJ0RFOFTDZtH-suFr3f5Wft |  | NA | NA | NA | NA | Good | Good | Easy to do registration | System maintained accuracy consistently throughout the surgery. | Pre-operative CT was used and registration done successfully later surgeon performed the skin sliding test , confirmed that the accuracy was perfect and marked craniotomy area. | Optical
2025-08-19 21:11 | hemarshitha@rxoom.com | Hemarshitha H. S. | Lakeview hospital | Belgaum | NS TRL 02 MF 204 |  | 1.2.24 | NA | https://drive.google.com/open?id=1O1K4cZ6hLeforRcVT10CA2sTGDskubpe | 2025-08-19 | Hemarshitha H S | Dr Raviraj Gorpade | X-Ray Spine |  | 15 | 0 | 0 | 0 |  |  |  |  | Yes |  |  | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | - | - | - | Improve the technique related to x ray based navigation | During the D6-D12 fusion surgery, the patient was initially positioned in the prone position. The surgeon recommended using X-ray-based navigation for the procedure. However, when attempting to fix the reference frame, it was noted that the area of interest was not adequately covered. To resolve this issue, I suggested fixing the location block on both ends to ensure proper coverage of the area. Unfortunately, the surgeon denied this request and decided to proceed with the case without navigation. | Optical
2025-08-26 16:00 | gajananbitode606@gmail.com | Gajanan Bitode | Nobel hospital | Pune | NS TRL 02 MF 201 |  | 1.2.24 | NA |  | 2025-08-26 | Gajanan manik bitode | Dr chandrashekhar | Cranial |  |  | 0.14 | 5.34 | 0.93 |  |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1T6rYLLkjEw6meWNPD6VzqeCcy5dI-PLr |  |  | Na | Yes hard drive not working due to someone inserted two cd | Na | Na | Na | Na | Na | Na | Na | Optical
2025-08-30 20:24 | gurjit758.232@gmail.com | Gurjit Singh | Care Hospital | Hyderabad | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1bK9-AapDh7aAN08j1VP5WEF6fUvorTK2 | 2025-08-30 | Gurjit Singh | Dr. Randhir Kumar | Cranial |  |  | 0.16 | 7.06 | 0.6 | https://drive.google.com/open?id=1vYbSKu1HJMCsQSlyQb9AR5f5EwtPQeBB, https://drive.google.com/open?id=1cJlwopv8L2Dyh85sGUQ_08B5izn7L7B6, https://drive.google.com/open?id=1UU6XAJYgA56T-4MWgajNOPGiWtiVY4uX |  |  | https://drive.google.com/open?id=1VOoOoUDMQBo9qUprUuuuSYHzsqhJLEdF | Yes |  |  | NA | NA | NA | Adding an instrument drawer/space to system to keep usual instruments in system. | By Randhir Kumar: Add VP shunt guide tool like Medtronic for VP shunt cases. | https://drive.google.com/open?id=1bMrl91I20YZ16V-_F9Tg43f3FoY61bfQ, https://drive.google.com/open?id=1vWope1YTmP-z8Slxl9xbqme_GH3nWtoN, https://drive.google.com/open?id=1FtxxkkG8bIa1bSm9eMxiQ_V7MQVmFkft |  |  | NA | NA | NA | NA | Strongly Agree | Strongly agree | Strongly agree | NA | Resignation done by assistant Surgeron and accuracy. | Optical
2025-09-01 22:15 | gurjit758.232@gmail.com | Gurjit Singh | Care Hospital | Hyderabad | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=19yhTBv4GutQRNIuMVIVkPTKyWiK4YTLm | 2025-09-01 | Gurjit Singh | Dr. Venkatesh Y and Dr. Bhuvneshwara Raju | Cranial |  |  | 0.16 | 4.75 | 0.68 | https://drive.google.com/open?id=1UAr5qQZ9dcShbj-MckMfx2qbEWVwe2gr |  |  |  | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=1oWSx206TAHhortIyFLUeGGb7QUStDsUC, https://drive.google.com/open?id=1yiqVDBnwbAXzucE-_IlirDh4l_Wqi-Jp, https://drive.google.com/open?id=1YTb7QZLP2oDFZlZxxxLIWxbV9Wygrzkw, https://drive.google.com/open?id=1ePSFJBc2oDzR8uawZQzyEMqHFoZGt00Z |  |  | Software crashed after completion of surface registration. | NA | Secondary MRI was not acquired according to navigation protocols | NA | Strongly Agree | Strongly Agree | Strongly Agree | NA | Patient position was lateral, ct scan was used as primary Dataset for model generation and patient registration, CT scan was taken in hospital and secondary Dataset 3 series 2D MRI(axial, coronal, sagittal) in which different views was clear due to 2D, MRI scan was acquired outside of the hospital without navigation protocols merged with ct scan on Dr. Request for better contrast view of MRI. On first patient registration by Dr. Venkatesh fiducial accuracy error was more than 5mm which I continued, on completion of surface registration software was crashed. Later I reloaded the treatment and deleted the merged series and did the registration and after completing registration then merged the MRI series again. Both Dr tested the accuracy and accepted the accuracy and completed the surgery successfully using navigation. | Optical
2025-09-01 22:37 | gurjit758.232@gmail.com | Gurjit Singh | Care Hospital | Hyderabad | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1_ybUupPbiA1583jOn3z5msZrOhs5-LpX | 2025-09-01 | Gurjit Singh | Dr. Venkatesh Yeddula and Dr. Bhuvneshwara Raju | Cranial |  |  | 0.16 | 4.75 | 0.68 | https://drive.google.com/open?id=1QEz07BoWdHsx-eWo50oE0sp8Yn9mXsFu |  |  | https://drive.google.com/open?id=1nsbKGQbU4OsBWjrFzAAqMQoc94zwZ9pq | Yes |  |  | NA | NA | NA | NA | NA | https://drive.google.com/open?id=15wpPZznk0ws37OO4mcQXthORQv_Cz_Rc, https://drive.google.com/open?id=12AFChz56l2TaQT6N-uPZr-No2mjR0KCG, https://drive.google.com/open?id=16DEquv64XlZsb53QeaIuZSEkr3ueQRrk |  |  | Software crashed after completion of surface registration. | NA | Secondary MRI was not acquired according to navigation protocols | NA | Strongly Agree | Strongly Agree | Strongly Agree | NA | Patient head position was lateral, ct scan was used as primary Dataset for model generation and patient registration, CT scan was taken in hospital and secondary Dataset 3 series 2D MRI(axial, coronal, sagittal) in which alternative view was clear due to 2D, MRI scan was acquired outside of the hospital without navigation protocols merged with ct scan on Dr. Request for better contrast view of MRI. On first patient registration by Dr. Venkatesh fiducial accuracy error was more than 5mm which I continued, on completion of surface registration software was crashed. Later I reloaded the treatment and deleted the merged series and did the registration and after completing registration then merged the MRI series again. Both Dr tested the accuracy and accepted the accuracy and completed the surgery successfully using navigation. | Optical
2025-09-03 00:23 | gurjit758.232@gmail.com | Gurjit Singh | Care Hospital | Hyderabad | NS TRL 02 MF 201 |  | 1.2.24 | NA | https://drive.google.com/open?id=1gaJ9ESw4D7XwgpHmYwTzHjrEDm31hUw4 | 2025-09-02 | Gurjit Singh | Dr. Venkateswara Yadulaa and Dr. Bhuvneshwara | CT-Spine |  | 1000 | 0.2 | 1.99 | 0.63 | https://drive.google.com/open?id=1VyG7oJAASqtw2h6JMK-jz23DZN7iA11p |  |  |  | No | Suppose to happen After single level L1 vertebra registration accuracy on that single vertebra should be accurate on navigation when navigation probe is placed on various points on L1 vertebra. Actually happened According to Dr. when Dr put probe on L1 Spinous process and lamina where fiducial were marked it show accurate but when he put probe on that same vertebra pedicle it probe show on one level below of that vertebrae when he put probe on L1 pedicle process it shows somewhere on L2 vertebrae and When he puts on D12 pedicle process it showed on L1 pedicle. | NA | NA | NA | NA | NA | Trackable spine instruments for pedicle screw placement. | https://drive.google.com/open?id=1cdBVdJnUC1oYPPxQC8F5dU37OXFvXQTh |  |  | NA | NA | NA | NA | NA | NA | NA | NA | Preoperative CT scan was used for spinal fusion. Dr decided to operate on L1 for scew fixation to use navigation, four fiducial points were marked on single L1 vertebra and point were crosschecked on 2D views and point correction was made. Dr did the registration got the acceptable accuracy and completed the registration. After completing the registration Dr. started checking the accuracy on spinal process and lower lamina where fiducials were collected were showing correct. According to Dr. he puts probe on L1 pedicle but on navigation probe show on L2. Same according him when he put probe on D12 pedicle on navigation probe show on L1 Pedicle. According to Dr. when he put probe on pedicle navigation show one level below difference while where were fiducial points where collected on L1 spinous process and lamina it is showing accurate. Re- Registration was done after marking again the fiducial points on L1 but same one level below difference was reported by surgeon when checked up on L1 pedicle. Probe was showing on L1 while collecting surface points and after registration completion checking on lower lamina and upper lamina it showing on L1 which is correct vertebrae level. | Optical
2025-09-04 10:09 | trushalgulhane1547@gmail.com | Faiz Khan | Neuro Max phalke | Pune | Other | Na | 1.2.24 | Na |  | 2025-09-03 | Trushal, gajanan | DR UMESH PHALKE AND SUMIT | Cranial |  |  | 0.26 | 4.56 | 1.8 |  |  |  |  | Yes |  |  | NA | NA | Na | NA | na |  |  |  | Na | Na | Yes slice thickness was 5mm and number of slices were only 38 | Na | NA | Na | Na | Na | Na | Optical