## Form responses 1
Timestamp | Email address | Date of surgery | Name of the Hospital | City | Did you face any issues while moving the system? (Write NA, if no issues were faced). | Name of the Surgeon | Application Engineer(s) | Type of surgery | Please provide information about specific type (sub-type) of procedure under the broad category of surgery identified above. Though not mandatory, this information could be useful for better understanding of the procedure. | What is the other type of surgery (Only if you have indicated Other above.) | Total Number of X-Rays During Spine surgery (Approximate, if exact number is not known. Write 1000, if Not applicable, or, Unknown. You can SKIP this question if this was not a spine surgery.) | Was the surgery successful? (Not the usefulness of our device) | Why the surgery was not successful? What was exact description of the failure (What was supposed to happen and what actually happened?) (You can SKIP this question, if the surgery was successful.) | To what extent our avigation system was used during the surgery? | Please write the reason why the navigation was used only partially, or not not at all. Please write what was expected to happen and what exactly happened that made the navigation not completely usable. (You can SKIP this question if the navigation was successfully used during the surgery.) | Were there any Instruments Related Issues? (You can SKIP this question if there were no instrument related issues?) | Were there any software related issues? (Please provide details. You can SKIP this question, if there were no software related issues.) | Were there any Data Related Issues? (You can SKIP this question there were no data related issues?) | Were there any other issues? (You can SKIP this question there were no other issues.) | Please write all the points that are significant and can help in improving system and demo practices. | Please provide a summary report on surgery. Please add whatever you think is significant point and can help in improving system and demo practices. | What are the inputs/requests from the Surgeon and/or hospital staff to the company? (Write NA, if none is provided.) | Were there any software related issues? (You can SKIP this question if there were no software related issues.) | Please provide Feedback from Surgeon related to Accuracy. (Write NA if surgeon did not give any feedback related to accuracy.) | Give Surgeon's Feedback related to Ergonomics. (Write NA, if surgeon did not give any feedback related to ergonomics.) | What was Surgeon's Feedback related to Ease of use? (Write NA if there was no feedback related to Ease of use.) | Did Surgeon gave any Feedback related to Anything else? (Write NA, if there was no feedback related to anything else.) | What are your inputs/requests to design and developments, logistics, assembly, etc.? (Please write NA, if you do not have any inputs at this stage) | Comments From Dr. Nilesh, Dr. Sundar, Sameer, Sanjay. (Please put your name after comment) | Any loss/damage to instruments seen after the surgery? Provide Details in short. (Write NA, if not seen.) | Any other information? Other Application Engineer (If any), Any issues with image, Any peculiar conditions, ... Write NA if you do not have anything to share. | System Deployed by | System Base - Optical, or, EM | Please provide the software version | System Identity | What is the other identity of the system? (SKIP this question if you have NOT answered "Other" above.) | Extra instruments provided with the system (Please write NA, if no extra instruments were provided available) | Tool Tip Accuracy in mm (Specify 100.00 if not available) | Fudicial Registration in mm (Specify 100.00, if not available) | Surface Registration in mm (Specify 100.00 if not available) | Upload system photograph | Upload Surgery Photos with the Doctor | Please provide navigation Accuracy and other relevant Photographs from surgery | Upload Demo Request Form | Upload consent Form | Upload Photo of Surgeon's feedback. | Column 47 | Is this Demo at a customer location?
Demo Details 2025 -2026
2025-09-04 23:28 | krishnakumar@rxoom.com | 2025-09-04 | ROYAL CARE HOSPITAL | NEELAMBUR | No | Dr. Rajkumar and Dr. Raguraja prakash | KRISHNA KUMAR M | Cranial | Temporal craniotomy + tumor excision | NA |  | Yes | NA | Completely | NA | NA | NA | Yes, but its not our side of issue. They only provide without contrast MRI dataset. So while usage it looks like not that much good. | NA | NA | Hospital: Royal Care Hospital City: Coimbatore System: Optical Surgeon: Dr. Rajkumar and Dr. Raguraja Prakash Surgery: Tumor Excision Procedure Summary: The patient's head was positioned laterally, and an MRI scan without contrast, taken at the same hospital, was used as the primary dataset for model generation, patient registration, and surgery. During the initial patient registration by Dr. Raguraja Prakash, the fiducial accuracy error was more than 5mm due to incorrect marking of fiducial points. I re-did all the points and marked 4 points myself, which Dr. Raguraja Prakash then re-collected. This improved both fiducial and surface accuracy. However, after registration, surface accuracy was slightly worse than fiducial accuracy because insufficient points were marked around the insertion area during registration. To ensure better results during navigation, I switched from surface accuracy to fiducial accuracy using the "use" button. Both Dr. Rajkumar and Dr. Raguraja Prakash tested and accepted the accuracy, and the surgery was successfully completed using navigation. | NA | NA | Neutral | Neutral | Agree | Need more precision | NA |  | NA | NA | Krishna Kumar | Optical | v1.2.24 | NS TRL 02 MF 204 |  | NA | 0.38 | 3.45 | 1.49 | https://drive.google.com/open?id=1mKCGA4dRRsXdLusL3GOpdW-i17Oej5YZ | https://drive.google.com/open?id=1w77ZgBv45dDJYQC7sDT7u0S03GbWG7r0 | https://drive.google.com/open?id=1m3DGp-01Pw-YoXiNrHk6BnzdRBWlLzs0 | https://drive.google.com/open?id=1CI8t2RNxQqU5aVw3_wkV3SODcjxcZmUA |  | https://drive.google.com/open?id=1lO4MmcZYhHavLxjxkyQCJV_fOUOy0BLa
2025-09-06 20:26 | gurjit@rxoom.com | 2025-09-06 | Asian ENT Hospital | Hyderabad | System height/top was more than OT gate at Asian ENT hospital, needed to remove camera arm and then shift the system in OT then mount the camera arm. | Dr. Chava Anjaneyulu, Dr. Katta Neelima, Dr. Aeesha Tasneem | Gurjit Singh | ENT | Right Ear MRM | NA | 1000 | Yes |  | Partially |  |  |  | Data was not taken from chin to vertex rest was good. |  | NA | CT Scan was used for the surgery. And skull tracker with sticker was used. Accuracy tested and accepted by Dr. Aeesha Tasneem, Dr. Katta Neelima. Navigation was used mid surgery to confirm the accuracy and in rest of the case navigation was not used much, according to Dr. based on surgery navigation was not much used required. | NA |  | Agree | Agree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.3 | 3.09 | 0.62 | https://drive.google.com/open?id=1s_B7dacdqmBhKIeD5dAtONA5wzLyv9Pw | https://drive.google.com/open?id=1jkHy2TkPPg3zsN_CWkivIIcGQpLd7Q7s, https://drive.google.com/open?id=1Vp6UN9iANdfs8bfbYJsA1EZ9Fs4EqRHt | https://drive.google.com/open?id=1oCt8d9ORMXD8gKWp2iQD1MB7E36sP45X, https://drive.google.com/open?id=1o3bmd2nWQ4NBK3pIFEsTlXPfz_rla3hO, https://drive.google.com/open?id=11FdfG8Kks6dDqstNaQ2jYkU5biFrfmks, https://drive.google.com/open?id=1eb9nu4RhI3kmA1B-oYVfhoD1Y0IbIWi6 |  |  | https://drive.google.com/open?id=1-fGFMhUJuzNEhtveXbdGPjaqU4CTiuIQ
2025-09-08 22:24 | gurjit@rxoom.com | 2025-09-08 | Care Hospital | Hyderabad | NA | Dr. T. Narasimha Rao, Dr. Manik Prabhu, Dr. Ajit | Gurjit Singh | Cranial | Tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was used for the navigation. Dr. T. N Rao did the accuracy test and accepted the accuracy. | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.13 | 4.82 | 0.9 | https://drive.google.com/open?id=1P4vp7j4pdxZsOk64SL7Gg_1TVG_bztAx | https://drive.google.com/open?id=173bajVh-elm286KWRLw1nADXQffHBoGO, https://drive.google.com/open?id=12CAMI7QfIwVV0rM8b3RSvVGLuUjwWdbI | https://drive.google.com/open?id=1pWIyX8B_7yQdvi9DB3_aavSuW-C4vDCW |  |  | https://drive.google.com/open?id=1gm35Qv2c85ZzOQ7QxZf_4j2i3Pj9J20a
2025-09-09 20:19 | gurjit@rxoom.com | 2025-09-09 | Care Hospital | Hyderabad | NA | Dr. T. Narasimha Rao, Dr. Manik Prabhu, Dr. Jersun | Gurjit Singh | CT-Spine | Spine Fixation L1-S1 |  | 40 | Yes |  | Partially | [5 empty] | NA | Per operative CT spine was loaded and used the crop feature to crop out the non operative vertebral area volume. Dr. did the registration on L4 vertebrae and achieved acceptable fiducial accuracy and completed the registration. Dr cross checked accuracy on left pedicle using navigation probe and C arm and then inserted screw in pedicle, after left pedicle screw insertion when accuracy checked on right pedicle accuracy was gone due to spine movement and re-registration was suggested but dr. didn’t go for it. When Dr move to L5 he did the registration but could not achieve acceptable fiducial accuracy and accuracy error was high on L5 then dr didn’t do redo registration. And complete the surgery with C-arm. | Design small size and steady spine reference attachment. |  | Strongly Agree | Strongly Agree | Strongly Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.35 | 4.24 | 1.8 | https://drive.google.com/open?id=1AwatV8ixL1alE0QFsc-B1a-f6fGSFMtH | https://drive.google.com/open?id=1olkBWkVWNeCHgJHV8-A2I1MH6OWs7naN | https://drive.google.com/open?id=15mGrC_uCn-UTH85SwJnDHO44eLhi3FOA, https://drive.google.com/open?id=1ndsyeSXfunqgm2MWECJdZ7Qebgb3avax, https://drive.google.com/open?id=1TNYzJ04RSODf-Llf56VeANzZtCHSpCPH |  |  | https://drive.google.com/open?id=1VxFaKLIriFAofeqlqiej4ib-XLP7brC1
2025-09-10 16:49 | trushal@rxoom.com | 2025-09-09 | Neuro Max phalke | Pune | NA | Dr. Umesh Phalke | Trushal Gulhane and gajanan bitode | Cranial | NA | NA | 0 | Yes |  | Completely | NA |  |  |  |  | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V2.24.2 | Other | Phalke hospital | NA | 0.4 | 4 | 0.8 | https://drive.google.com/open?id=1YLnJU8zD8FMVWrKXhhOiU60V4eQMNg1r | https://drive.google.com/open?id=1JlsG_iE-3Qtmoz822rcE1Sy2ywvuypo6 | https://drive.google.com/open?id=1d15bnItU171lp-u5SYWik2PEGK-2_k4D
2025-09-10 21:09 | gurjit@rxoom.com | 2025-09-10 | Care Hospital | Hyderabad | NA | Dr. T Narasimha Rao, Dr. Ajith , Dr. Rahul | Gurjit Singh | CT-Spine | Spine Fixation L4-L5 |  | 30 | Yes |  | Partially | [5 empty] | NA | Per operative CT spine was used. Dr. T N Rao did the registration on L4 vertebrae and achieved acceptable fiducial accuracy and completed the registration. Dr cross checked accuracy using navigation probe and C arm on L4 and said accuracy is okay after operating on L4 when Dr start operating on L5 without re- registration accuracy was off on L5 and Dr didn’t go for re registration of L5 and choose to complete the surgery with C-arm only. | NA |  | NA | NA | NA | NA | NA |  | One passive marker coating. | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.22 | 4.01 | 0.67 | https://drive.google.com/open?id=1KaQHJgcM0832ntFCYT7nF3f3AkWkXt-y | https://drive.google.com/open?id=1YTuqvYEpIOIiEWrx_DGBjDcvxYS_vPOi | https://drive.google.com/open?id=1APTqDz-LI8VoDAiOzHqSA0ky2P1c4kak, https://drive.google.com/open?id=14zHARfQDTOZFqU1Z5mKI_H57-1dbV4u2
2025-09-12 20:00 | gurjit@rxoom.com | 2025-09-12 | Asian ENT hospital | Hyderabad | NA | Dr. Chava Anjaneyulu and Dr. Aeesha Tasneem | Gurjit Singh | ENT | Seploplasty + FESS |  |  | Yes |  | Partially |  |  |  | Vertex was not covered. |  | NA | Preoperative CT scan was used and skull tracker with sticker used and registrations was done two time due to some accuracy error observed despite achieving acceptable accuracy in fiducial and surface but Dr. accepted the accuracy. After completing the surgery Dr. used navigation and said accuracy is workable but they need 70 degree curved probe for frontal and maxillary sinus. And Suction probe or tool for calibrating their suction So they can do surgery seamlessly otherwise they keep switching between suction and probe while using endoscope. | Required 70 degree curved probe for frontal and maxillary sinus, any tool that can calibrate with their suction, suction probe- straight and curved. |  | NA | NA | NA | NA | Different angled long curved probe, software workflow and tool to calibrate any surgical tool during surgery. |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.13 | 3.93 | 0.44 | https://drive.google.com/open?id=1_hliXiKUPKBh_3pcHUIexQAEJ8ZypVzt | https://drive.google.com/open?id=1p2LJe45R3MszTcq4-u8P6CwFEcZv2MJL | https://drive.google.com/open?id=1I5JvG3uyRDXCymltkxCJpfrCaa65wquT, https://drive.google.com/open?id=1zmFGFLrp5e_4xMwLL-JDIBondbMh0buA, https://drive.google.com/open?id=1fgaZzzgOa_huFdX-s1S7F5x2d2mEMbzQ |  |  | https://drive.google.com/open?id=13rv1HIsUKjlb2lkevfH3tCsMSkeUWeFm
2025-09-13 21:52 | gurjit@rxoom.com | 2025-09-13 | Care Hospital | Hyderabad | NA | Dr. Randhir Kumar | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Tumor Biopsy |  |  | Yes |  | Completely |  |  | During the Fiducial registration, after capturing the first fiducial mark on patient and when I go for capturing second Fiducial point it was not registering the second point, no sound only processing bar showing of registering the point, but on completion of processing just zooming effect were happening and Fiducial point was not registering. |  |  | NA | Preoperative CT scan, skull tracker with mini plate base was used and trajectory planning was done by Dr. While Patient registration a software issue occurs that second fiducial was not registering instead zooming effect was happening. Only creating new treatment solved the point not registration error, Dr. marked the trajectory again, registration was done and Dr tested and accepted the accuracy. With probe marking was made and using needle guidance tool Dr took the biopsies. | To add trackable biopsy needle also. |  | NA | NA | NA | NA | Provide a needle measurement scale like Medtronic so Dr can easily apply the stopper at right distance on any biopsy needle. |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.24 | 3.85 | 0.54 | https://drive.google.com/open?id=1Om6ID1oFqdc9ee0XiM-AWW-2PdcXQPOT | https://drive.google.com/open?id=1qf1zmxAYRewmdCFaRTHETwJuPJ41ihGM, https://drive.google.com/open?id=1Vumm_OSmq6bQHVYKyI2WGxjTQX44ezqY, https://drive.google.com/open?id=1pMqIt7J_lgYKwnc-TCC6sqHAfjHcrlkh | https://drive.google.com/open?id=1KnDofCJIM44Ndu5LaLyBdH8oJfQahQbg, https://drive.google.com/open?id=1XHNWxpx-JLQYxAYs9OIy0a0mYaMq07pn, https://drive.google.com/open?id=1pULAx9A3U54rqHstSUpUQLC5TlrCamic, https://drive.google.com/open?id=1HqtWuqc8Pd61bRYv29lAMERpWGdrYMnt, https://drive.google.com/open?id=16qjSZVW2w7i6Oiw884mWP4eyS7XhvvEl
2025-09-16 23:48 | gurjit@rxoom.com | 2025-09-16 | Care Hospital | Hyderabad | NA | Dr. T N Rao, Dr. Manik Prabhu, Dr. Jersun | Gurjit Singh | Cranial | Tumour Excision |  |  | Yes |  | Partially | Consistent accuracy during the surgery was supposed to happened. But in accuracy deviation found after the incision was made. May due to movement in Reference or head clamp Sugita. |  |  |  |  | NA | Pre created treatment was loaded in which CT scan was used as primary dataset and MRI was merged. Registration was done and accuracy was tested and accepted by Dr. Accuracy was checked after the draping which was accurate then accuracy was checked after incision m in which some accuracy deviation was observed. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.33 | 1.43 | 0.52 | https://drive.google.com/open?id=1flsjfyQVdEdK_ODSq1kxsmaRdEij70Kt | https://drive.google.com/open?id=146pr4TzGsNJDt76daycEmidvTKUky0or | https://drive.google.com/open?id=1oGE4suU95sRMhr2W1BenXG52CLXoifnb, https://drive.google.com/open?id=1LQGYGCGYxo3wSWZAC4OxCqaRRW0Qtqt5
2025-09-19 16:06 | gajanan@rxoom.com | 2025-09-17 | Apollo sage hospital | Bhopal | Na | Dr Rahul jain sir | Gajanan bitode | Cranial | Pituitary tumor excision | Pituitary tumor excision |  | Yes |  | Completely | [5 empty] | No suggestion from doctor he likes our system | According to Dr Rahul jain sir machine accuracy was very good and also taking less place as compared to Medtronic | Doctor needs some more demos in upcoming days |  | Rahul jain like accuracy | Comparing system with other and he finds out system is less bulky | NA | NA | NA |  | NA | Na | Gajanan Bitode | Optical | 1.2.24 | NS TRL 02 MF 201 |  | Na | 0.29 | 4.21 | 0.98 | https://drive.google.com/open?id=1_EGWaEOsrIZwoGun6KrH7FpiKtSOMYSj | https://drive.google.com/open?id=1_fAo8A2PU096_aTIwwxV_gINZFkUyTbV, https://drive.google.com/open?id=1Zdd_xTaeJdHqjF6nl70ndl3RT2s-rVem | https://drive.google.com/open?id=14wcTs3rBQPVig1HmlQZyguaOsbLVT3D6
2025-09-19 17:25 | mrkrishnakumar.2105@gmail.com | 2025-09-19 | KMCH | Coimbatore | NA | Dr. D. Ganesan | Krishna kumar M | Cranial | Right craniotomy + tumor excision |  |  | Yes |  | Partially | Due to the location of the tumor in the patient's head. He used our navigation system to check whether the anatomical landmarks were correct or not. Also, he used for getting clarified results after removal. |  |  |  |  | NA | Case Report - Hospital: KMCH Hospital - City: Coimbatore - System: Optical - Surgeon: Dr. D. GANESAN - Surgery: Tumor Excision in right craniotomy Procedure Summary: The patient's head was positioned Supine, and an MRI scan taken at the same hospital, was used as the primary dataset for model generation, patient registration, and surgery. During the initial patient registration. All the points and marked 5 points myself, which Dr. D. Ganesan then check those points whether are correctly pointed specifical places or not. Then I got good results of accuracy after I done registration. After registration, surface accuracy and fiducial accuracy was good because of i marked all the points and recollect all the points and covered around the insertion area during registration. To ensure better results during navigation, I explained switching feature from surface accuracy to fiducial accuracy using the "use" button. Dr. D. Ganesan tested and accepted the accuracy, and the surgery was successfully completed using navigation. The tumor was big in size and Dr. D. Ganesan operate OT out and also used out system effectively. Accuracy: - Fiducial accuracy: 2.72 mm - Surface accuracy: 0.79 mm Surgical Outcome: Successful | NA |  | NA | Very compact and ease to place in OT also. | Registration Feature accessing aren't complex way | NA | NA |  | NA | NA | Krishna Kumar | Optical | v1.2.24 | NS TRL 02 MF 204 |  | NA | 0.13 | 2.72 | 0.79 | https://drive.google.com/open?id=11Qqy3Bw9pilV8pSh0oY56ouSWDzl2Otm | https://drive.google.com/open?id=17RYkIU0M1ma2BFtYyyidhSQGMDz1Gr-M, https://drive.google.com/open?id=1s_NsdVbgnPslfw8QUFd7Dvb_BDUJhWKc | https://drive.google.com/open?id=1tRniAXpPtFCiCWRzQDunWYIUpD2sn36q | https://drive.google.com/open?id=1u2lYbZLWLow5l9gB_UhcKhd5q9Q0-RwH |  | https://drive.google.com/open?id=1YkFFFPSmUiPvOAHYmywQYlSQqeaKtraA
2025-09-23 22:13 | gurjit@rxoom.com | 2025-09-23 | Care Hospital | Hyderabad | NA | Dr. Randhir Kumar | Gurjit Singh | Cranial | Pituitary Tumor Excision |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative MRI was used as primary data set and CT scan was merged with it. Skull tracker with mini plate was used. Registration was done and accuracy was tested and accepted by surgeon. | Develop a tool to calibrate and track endoscope. Develop a better spine reference attachment which can be screw/drill in spine something like K Wire. |  | Agree | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.35 | 3.66 | 0.53 | https://drive.google.com/open?id=1C1HSCBcTomvmyqBwyFax0FpDK17eOTIC | https://drive.google.com/open?id=1sUFgbXutmUMYsqal1E0pE-qwdPhhp_oj, https://drive.google.com/open?id=1fJE7U2ASFhHeb-kzb5GUMMf8BYaR9stt | https://drive.google.com/open?id=1Njmvlc0vHmU0JlRlHAnMt28aXqRfFFtn, https://drive.google.com/open?id=1CVOwMJKwGYnc6GRGgXh9Sf9FkvYrElRp
2025-09-23 23:22 | rohan@rxoom.com | 2025-09-20 | Techno India Dama Hospital | Kolkata | Yes, Upper acrylic cover has been broken | Dr. Amitava Das | Rohan Ghosh | Cranial | Tumor Excision | NA | 0 | Yes | NA | Completely | NA | NA | NA | NA | NA | Clinical Navigation Report Hospital: Techno India Dama Healthcare and Medical Center, Kolkata Date of Surgery: 20th September 2025 Surgeon: Dr. Amitabha Das Surgery: Tumor Excision Surface Registration Accuracy: 0.86 mm CT contrast data was used for navigation. Dr. Amitabha Das performed the registration himself. On the first attempt, wrinkles on the elderly patient’s skin caused the probe to slip, leading to an incomplete registration. On the second attempt, registration was successfully completed, though fiducial accuracy exceeded 5 mm. Despite this, the surgeon confirmed accuracy remained intact throughout surgery and expressed overall satisfaction. Feedback highlighted that every feature of the system is very useful, with no negative observations. | CT contrast data was used for navigation. Dr. Amitabha Das performed the registration himself. On the first attempt, wrinkles on the elderly patient’s skin caused the probe to slip, leading to an incomplete registration. On the second attempt, registration was successfully completed, though fiducial accuracy exceeded 5 mm. Despite this, the surgeon confirmed accuracy remained intact throughout surgery and expressed overall satisfaction. Feedback highlighted that every feature of the system is very useful, with no negative observations. | Dr. Amitabha Das personally carried out the registration process. On the first attempt, due to the patient’s advanced age and wrinkled skin, the probe slipped from its position, resulting in an incomplete registration. On the second attempt, registration was completed successfully, though the fiducial accuracy exceeded 5 mm. Despite this, the surgeon reported that navigation accuracy remained intact throughout the entire procedure and did not affect the surgical workflow. The surgeon highlighted that all features of the NeuroSaphireX system were very useful, with no negative feedback provided. He strongly agreed that the system was easy to load patient data, the registration process was simple, navigation was accurate, instruments were ergonomic, and the user interface was simple. Overall, Dr. Amitabha Das expressed satisfaction with the system’s performance and usability, appreciating its reliability and supportive role during surgery. | NA | Every feature is useful | Every feature is useful | Every feature is useful | Every feature is useful | NA |  | NA | NA | Rohan Ghosh | Optical | 1.2.24 | NS TRL 03 SE 101 |  | NA | 0.08 | 7.31 | 0.86 | https://drive.google.com/open?id=1QllHGfoCtsgJ8ya70YDCdWlHIs1MsTIz | https://drive.google.com/open?id=1TAss9z0StPqWEpfwEkJUqB_RW-BU3_64, https://drive.google.com/open?id=1eDb7OiVlWC5_vii1_jhhqau-iYXDCMjT | https://drive.google.com/open?id=1fPIEMXytry6L7KufrLEdTPFukfr0t0O- | https://drive.google.com/open?id=19JcpoGIZFAIcpuHV5VAch9rPgO0Uduh0 |  | https://drive.google.com/open?id=1YsJ3HXInUIPK1LepnvwIcntEj7Rttw9l
2025-10-01 17:32 | gurjit@rxoom.com | 2025-09-29 | Asian ENT hospital | Hyderabad | NA | Dr Chava Anjaneyulu, Dr Aeesha, Dr Neelima | Gurjit Singh | ENT | Transnasal Endoscopic |  |  | Yes |  | Partially |  |  |  |  | Patient Data first cd received was not readable. Then second CD came in mid of surgery | NA | First patient data DVD was received had issue, it was not getting read by any system. Scan was taken out side of the hospital. Then Dr Chava asked the CT scan center to resend the DVD. In mid of surgery CD arrived then Dr asked us to do the navigation to at least check the accuracy. Dr wanted to use sticker base skull tracker but in the end standard reference with arm decided to use. Registration was done and confirms by Dr. Aeesha. Dr Chava also confirmed the accuracy. | NA |  | Agree | Agree | Agree | Surgeon asked the virtual probe tip to be smaller. | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.35 | 4.78 | 0.49 | https://drive.google.com/open?id=1tl2Us3vePoVl_54OM4l4OAGao73Z64GE | https://drive.google.com/open?id=1XfyMEGOQTYTtoZ2GGFlpPp_DSwfddF6S, https://drive.google.com/open?id=1k9accO0tD1fJmDNZCNOotAWZaJBic3Xr | https://drive.google.com/open?id=1SFcnqsO03Nz-_7qNF_-fTtcArtvo9Ml1, https://drive.google.com/open?id=19ciMrMBSkobKXF98KKkLCzBYE5y3Hjvz, https://drive.google.com/open?id=1s7Ba5kpZIdnETc86gtEmhf5lb3QHWE4s, https://drive.google.com/open?id=1IjSThCCPqDMmsHm7Dtxc5xOey1eWuxul, https://drive.google.com/open?id=12tw7wJ6_LwKuGd7-y8xKnzLSoO6r5K03 |  |  | https://drive.google.com/open?id=107TP-4t-2qUZZV-8_eD9jhNECvqbeuYj
2025-10-07 14:40 | trushalgulhane1547@gmail.com | 2025-10-06 | Hinduja hospital | Mumbai |  | Dr. Sameer bhargav | Trushal Gulhane | ENT |  |  |  | No | It was noted that while accuracy was within acceptable limits before draping, a deviation occurred afterward—possibly due to reference movement during draping, which could have affected tracking precision | Partially | It was noted that while accuracy was within acceptable limits before draping, a deviation occurred afterward—possibly due to reference movement during draping, which could have affected tracking precision |  |  |  |  | We need to emphasize on the use of skrew of skull tracker to obtain good accuracy throughout the case | Pre-operative Phase: The pre-operative CT scan was utilized as the primary dataset for image guidance. Skull tracking was performed using a mini plate, secured with Dynaplast and reference stickers. System registration was successfully completed, and accuracy was verified and accepted by the operating surgeon prior to draping. Intra-operative Phase: Following draping, tool-tip calibration was performed successfully, and a long probe was used. During navigation, the probe’s appeared inside the orbital region, indicating a deviation. I requested the surgeon to perform re-registration to restore accuracy; however, the surgeon elected to proceed without navigation assistance. Post-operative Discussion: After completion of the surgery, a discussion was held with the surgeon. It was noted that while accuracy was within acceptable limits before draping, a deviation occurred afterward—possibly due to reference movement during draping, which could have affected tracking precision. The surgeon acknowledged this observation. Surgical Usage: Procedure completed successfully. Navigation system used for initial registration and accuracy verification. No intra-operative navigation guidance was utilized following observed accuracy deviation. | NA | NA | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.2.24 | Other | Hinduja system | NA | 0.2 | 3 | 0.6 | https://drive.google.com/open?id=1EUpVNUyYD6NY5QgZdFaUr3vLaw615iOd | https://drive.google.com/open?id=1ouScspY_GArOL2rBEy-IylRqDvGo3Fl5 | https://drive.google.com/open?id=1ehyD7Xfz9cOUwMGM_dnO3GdOzlGMnjkR
2025-10-07 19:24 | gurjit@rxoom.com | 2025-10-06 | Care Hospital | Hyderabad | NA | Dr. Venkateshwar Yadulla and Dr. Bhubneshwar | Gurjit Singh | Cranial | Craniotomy & Tumor Excision |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan used as a primary data and MRI scan was merged with it. Registration was done accuracy was tested and accepted by Surgeon. Surgery completed using the navigation. | NA |  | Good accuracy. | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.07 | 2.78 | 0.39 | https://drive.google.com/open?id=13D1-L66U1nQE2GHt_bVo_moyoB33B68W | https://drive.google.com/open?id=166AC9X0E_Eel7IUG5dMlIzd_MR2sbQA9 | https://drive.google.com/open?id=109VJhUKbZnwel82Jap1g-3jlEa2bmjfc, https://drive.google.com/open?id=1usz1ijjEC-I8rDFxAL428twhaJ4oiFAb
2025-10-07 21:03 | gurjit@rxoom.com | 2025-10-07 | Care Hospital | Hyderabad | NA | Dr. T N Rao, Dr. Bhuvaneswara Raju, Dr Jersun | Gurjit Singh | Cranial | Craniotomy &Tumor Excision |  |  | Yes |  | Partially |  |  |  | Data coverage was not as per protocols; coverage chin to vertex |  | NA | Pre operative CT scan used as a primary data and MRI scan was merged with it. First Registration was done by Dr. Jersun and accuracy was tested by Dr. T N Rao and said there is 2-3 mm accuracy error on the mid line of brain then re registration was done again by Dr Jersun. Dr still said the same error is there and I captured the difference. Then I did the registration again accuracy was better than before but Dr T N Rao said accuracy error is still 2-3 mm on the mid line before draping and after completing the surgery. Navigation was used to confirm the accuracy. | Improve the accuracy. |  | Okay but can be improved | NA | NA | NA | NA |  | One glyonce coating was removed. | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.15 | 2.68 | 0.56 | https://drive.google.com/open?id=1Bhn-AXb9REaJYKae9fzSeObu2XzEFKbj | https://drive.google.com/open?id=1wrmb_SQ7GSksUwhehHHqd-3S0kDFYoLf, https://drive.google.com/open?id=1LFEyBH3Ty9XMoO_OJ4kIXDR5ih2GdhF9 | https://drive.google.com/open?id=1A2NgS7vz1oNUK6_b0Cl-Tt9aiidTdiNM, https://drive.google.com/open?id=1TN8l3nwvVxr-XKT5YZwDECvdvjKvKO2t, https://drive.google.com/open?id=1nKlqJ8WJi1UT1cpbntouQJz1fBj_EPF1, https://drive.google.com/open?id=1t02-x15tva5FAgCOT8oWqMbeKE9n754U, https://drive.google.com/open?id=1PCeezIyjtHqMTqNlw2t6CezZ74pSVv5p
2025-10-07 21:42 | gurjit@rxoom.com | 2025-10-07 | Care Hospital | Hyderabad | NA | Dr. Manik Prabhu and Dr. Rahul | Gurjit Singh | Cranial | Craniotomy & Tumor excision |  |  | Yes |  | Not used | Suppose to happen: After patient registration and accuracy should be good on patient. What happened: Registration was done multiple times but accuracy error was high on patient. |  |  |  |  | NA | Pre-operative CT scan was loaded and skull tracker with mini plate was decided to use. Skull tracker with mini plate was applied and multiple registration attempts were made to get good accuracy and upon checking accuracy the accuracy error was high and Dr. decided to not to use navigation | NA |  | NA | NA | NA | NA | NA |  | NA | Patient was on ventilation while scan was taken which was captured in scan images as well. | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 100 | 4.8 | 0.53 | https://drive.google.com/open?id=1SiT4hUO-TiAwWrziOb_-PCdgVo4yVBQd | https://drive.google.com/open?id=1b-wN7LMHk9KzbVopwoAlnTlEBgD1BWbP | https://drive.google.com/open?id=1RBbK3vpVLasLsgetrZcrHy-suf41CHLV
2025-10-09 15:58 | mrkrishnakumar.2105@gmail.com | 2025-10-09 | KMCH | Coimbatore | NA | Dr. D. Ganesan | KRISHNA KUMAR M | Cranial |  |  |  | Yes |  | Partially | Because he needed to confirm the location of the tumor before making anatomical markings and after opening the skull. |  |  |  |  | NA | The patient's head was positioned supine, and an MRI scan taken at the same hospital with contrast had been used as the primary dataset for model generation, patient registration, and surgery. During the initial patient registration, I had marked all 5 points and re-collected them, which Dr. D. Ganesan then checked for correct placement. I had achieved good results for accuracy after registration. Surface accuracy and fiducial accuracy were good, particularly around the insertion area. The surgeon used the S1 (short probe) for initial and intermittent checks to ensure the procedure was on track, given the tumor's size and proximity to the skull. Although two cases were scheduled for today, the tumor resection took longer than expected due to the tumor's size and proximity to bulged veins/arteries. | NA |  | He's strongly agreed | He's strongly agreed | He's strongly agreed | NA | NA |  | NA | NA | Krishna Kumar | Optical | V1.2.24 | NS TRL 02 MF 204 |  | NA | 0.54 | 3.75 | 0.43 | https://drive.google.com/open?id=1v6H1MxdafKMLjcXz-wfuMcahMyPm2w0H | https://drive.google.com/open?id=1f1EV6c-qi4rWAIWm2jekOuCiHVdWBFQc | https://drive.google.com/open?id=1RkHOuQ7B4E85pOdSBXxu5uZ7td1Jd3UM | https://drive.google.com/open?id=1L6o_LBCVyCwjycgCZU52TcjlNnYCzzPk |  | https://drive.google.com/open?id=1E2PrVPnowkwubEdXW7YGPMwjiu7NP7Pt
2025-10-10 01:53 | mrkrishnakumar.2105@gmail.com | 2025-10-04 | KMCH | Coimbatore | NA | Dr. D. Ganesan | KRISHNA KUMAR | Cranial |  |  |  | Yes |  | Not used |  |  |  |  | This is accidental, a manual error occurred after the dropping( Redo button press issue after s2 stylus calibration ).So the process could not continue, but surgery done successfully without help of our system. | NA | Case Report - *Hospital:* KMCH Hospital - *City:* Coimbatore - *System:* Optical - *Surgeon:* Dr. D. Ganesan - *Surgery:* Tumor Excision *Procedure Summary:* The patient's head was positioned supine, and an MRI scan taken at the same hospital had been used as the primary dataset for model generation, patient registration, and surgery. During the initial patient registration, I had marked all 5 points and re-collected them, which Dr. D. Ganesan then checked for correct placement. I had achieved good results for accuracy after registration. Post-registration, surface accuracy and fiducial accuracy were good, particularly around the insertion area. However, during the tool tip calibration with the S2 probe (long probe), an error had occurred where the left side of the column displaying accuracy details for tool tip accuracy, fiducial registration, and surface registration vanished. Despite restarting the system and rebooting, the issue had persisted. *Discussion:* After discussing the issue with our team (Anita ma'am, Hemarshita ma'am, and Manoj sir), I understood the reason behind the error. Dr. D. Ganesan had been initially impressed with the precision and happy with the registration accuracy. The error that had occurred later was the main concern for the demo. For future cases, I will be more cautious and take steps to avoid similar issues. *Initial Accuracy Achieved:* - Fiducial accuracy: 4.00 mm - Surface accuracy: 0.68 mm *Surgical Outcome:* The surgery was successful without our navigation system. *Doctor's Feedback:* Dr. D. Ganesan had provided positive feedback on the initial results, but the error that occurred later was a concern. Despite this, he was understanding, given the good initial accuracy. | NA |  | Initially it was good. But later, the s2 calibration it gone wrong. | Strongly accepted | Strongly accepted | he also want to know...what was happened really... | NA |  | NA | NA | Krishna Kumar | Optical | V.1.2.24 | NS TRL 02 MF 204 |  | NA | 0.038 | 4 | 0.68 | https://drive.google.com/open?id=1xNznqUGXTGwqAEm-IGUwAh5-ax-IxqSv | https://drive.google.com/open?id=1Xvy9ik-szTQB103-0fHENHnIuJePnS8G | https://drive.google.com/open?id=19llZYOfEUOXndqvYJuQ4HAasDTNNZ65b
2025-10-11 10:26 | gurjit@rxoom.com | 2025-10-10 | Star Hospital | Hyderabad | NA | Dr Ajay Reddy | Gurjit Singh | Cranial | Pituitary tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was loaded and used as primary dataset and pre operative MRI was merged with it. Skull tracker with minimal plate with screws applied on the patient head and registration was done by Surgeon and accuracy confirmed and accepted by surgeon. Surgery successfully completed using navigation. | NA |  | Agree | Strongly Agree | Strongly agree | Dr expected EM to use to don’t have line of sight of issue. So he had some issue to maintain line of sight during surgery like unknowingly covering the glyons. | NA |  | One Glyons coating | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.36 | 3.2 | 0.53 | https://drive.google.com/open?id=1pC_svbdh_BQ8Ih7cT_v-OIClfpxcdZqa | https://drive.google.com/open?id=1fWgSckdFJukXs4afFY2djbDXBngoFFHn, https://drive.google.com/open?id=1cRAg8FWzl6z9h2cCGvJS5Hy2_fNOE4TY | https://drive.google.com/open?id=1xr_Wwo-bpA4cYPmTKIaEMxeoVUX7hs-o, https://drive.google.com/open?id=1na_hCMU8Ozfm2PhoW-2YMvrVHpu49HSe, https://drive.google.com/open?id=1LMlpI_EVobCH8ZNWKkhcWDREvwyEksXR | https://drive.google.com/open?id=1PIxJMD584NV2j_rX3C3qt9whZAfswyrP |  | https://drive.google.com/open?id=1RyRdXetwfQms27W4hSAtHPIGyHMGe1i_
2025-10-21 20:29 | gurjit@rxoom.com | 2025-10-21 | Osmania General Hospital | Hyderabad | Yes while unloading there was no man power, they were stuck place due to rainin somewhere. | Dr. Mastan Reddy, Dr. Srikant | Gurjit Singh | Cranial | CP angle tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative ct scan was loaded as primary data and secondary MRI data was merged and registration was done and accuracy tested and accepted by surgeon. Surgery completed using navigation. | Tool to Use any surgical tool as navigable instrument. |  | Agree | Agree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.17 | 3.68 | 0.45 | https://drive.google.com/open?id=1AgdEaKiUQ9U2_M55J_4hVKIdHz77PeAl | https://drive.google.com/open?id=1q0DQYyRUn2CXvybVi59Vp-T_Q2frO6L8 | https://drive.google.com/open?id=1Ex5GbAaxZNief305pzrvPiHanIh1W_9d | https://drive.google.com/open?id=1n3sONRYKeOQ5XBye1MBwfZAH8Y0OXbqY |  | https://drive.google.com/open?id=1Zzo_Y9kls6ki7XNDYeEMZXgk5_QKGq2u
2025-10-22 20:12 | gurjit@rxoom.com | 2025-10-22 | Sindhu Hospital | Hyderabad | NA | Dr. Narasinga Rao, Dr. Anish | Gurjit Singh | Cranial | VP Shunt |  |  | Yes |  | Completely | [5 empty] | NA | The pre-operative CT scan was used as the primary dataset, and the MRI scan was merged as secondary data. Dr decided to use Skull tracker with sticker. Skull tracker was applied. Registration was performed successfully, and accuracy was verified and accepted by the surgeon. Dr confirmed the trajectory with probe then Dr tried to use VP Shunt tool with guidance needle but Dr. was not much confident on it so Dr. tried to place the shunt without using VP shunt tool but did not hit the desired spot/ventricle then he change the trajectory and put again then hit the desired spot/ventricle and said VP shunt tool was correct. Navigation was used completely throughout the surgery for guidance and confirmation. | Make light weight skull tracker to use it better by just sticking on head using sticker and ioban. |  | Strongly agree | Disagree | Agree | Option to flip the views of patient scan. | NA |  | NA | Dr Said in in S8 Medtronic they have a light weight skull tracker which can be stick on head and they don’t required to use screws to fix the tracker. Late discussing with Dr that it may be EM that they have used. But Dr said it was optical based one. | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.39 | 3.14 | 0.47 | https://drive.google.com/open?id=1N5oL0NKp3KlpUO8I6JVXUKVTUKP0QzGM | https://drive.google.com/open?id=1Joq-O1eAbmSHqAbvrKsIYeLHergjgOeB | https://drive.google.com/open?id=1svzKtXQXp6_7TgNObHcsBHIVyS8pY-Yi, https://drive.google.com/open?id=1jHO5YziYyr8Y3g9wRDObBC-CB11awF7_ | https://drive.google.com/open?id=19iQyxMM7hxr7WiRdZnAOc-U-IRqLm4PA |  | https://drive.google.com/open?id=1htlfaRaeJql_Qk5TmHMf0BLMT_5gnjZk
2025-10-22 20:37 | gurjit@rxoom.com | 2025-10-22 | Sindhu Hospital | Hyderabad | NA | Dr. Narasinga Rao, Dr. Anish | Gurjit Singh | Cranial | Craniotomy tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | The pre-operative CT scan was used as the primary dataset, and the MRI scan was merged as secondary data. Registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA |  | Strongly Agree | Agree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.16 | 3.49 | 0.27 | https://drive.google.com/open?id=1J1gzB0Q38_uyDgcy2Pf3rjhj-NI_mI_X | https://drive.google.com/open?id=1HJiiztitlzQIikLyEM0oUp_PbluU9EE3, https://drive.google.com/open?id=1oOZSpLvYcuy9DbUdfYd0HwT7b_UPrEpC | https://drive.google.com/open?id=114AiqOzLODzlbK6Xu4f1uOHdZGhv2MG8
2025-10-25 23:43 | mrkrishnakumar.2105@gmail.com | 2025-10-25 | KMCH | Coimbatore | NA | Transnasal transphenoidal excision of pituitary adenoma | KRISHNA KUMAR | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | _Case Report_ _Hospital:_ KMCH Hospital _City:_ Coimbatore _System:_ Optical _Surgeon:_ Dr. JKBC Parthiban _Surgery:_ Transnasal transphenoidal excision of pituitary adenoma _Procedure Summary:_ The patient's head was positioned supine, and an MRI scan taken at the same hospital was used as the primary dataset for model generation, patient registration, and surgery. During the initial patient registration, I marked all 4 points and re-collected them, which Dr. Malini, junior surgeon of Dr. JKBC Parthiban, checked for correct placement. I achieved good results for accuracy after registration. Post-registration fiducial accuracy was good compared with surface registration, particularly around the insertion area. The surgeon continuously used our navigation. _Accuracy Achieved:_ - Fiducial accuracy: 4.72 mm - Surface accuracy: 0.96 mm _Surgical Outcome:_ Case successfully completed _Doctor's Feedback:_ Dr. JKBC Parthiban provided positive feedback, but OT dress unavailability led to a delay in entering the OT. I initially discussed the reason for my belated entry, and later he understood. | NA |  | _Doctor's Feedback:_ Dr. JKBC Parthiban provided positive feedback, but OT dress unavailability led to a delay in entering the OT. I initially discussed the reason for my belated entry, and later he understood. | Agreed | Agreed | NA | NA |  | NA | NA | Krishna Kumar | Optical | v1.2.24 | NS TRL 02 MF 204 |  | NA | 0.27 | 4.72 | 0.96 | https://drive.google.com/open?id=1viWyrSR71okrR_hnDVP_bDgUONLpOjtT | https://drive.google.com/open?id=16LV1ZNrGY9T0YTjs9iTPzM0fn5GaevMH | https://drive.google.com/open?id=1zNPnhVN3zfxk7X0gTQ6RADFAMOK7lNVl | https://drive.google.com/open?id=1mlnuxjvNk9NuHVtTOqkjkIW-6Rk5Vxn_ |  | https://drive.google.com/open?id=13iusk2Mxc7oEdc8_5IsyXmhgA8zRWxnz
2025-10-29 18:48 | trushalgulhane1547@gmail.com | 2025-10-29 | Kiran hospital | Surat | NA | Bhaumik thakor | Trushal Gulhane | Biopsy, Any (Cranial, Full body) | NA | NA | 0 | Yes |  | Completely | [5 empty] | NA | Surgery Report Date: 29 October 2025 Technology Used: Optical Navigation System Hospital: Kiral Hospital, Surat Surgeon: Dr. Bhaumik Thakor Case: Cranial Biopsy Accuracy Achieved: 0.5 mm Data Used: CT and MRI Procedure Summary: Pre-operative CT and MRI data were used for registration and planning. The optical navigation system was successfully set up, and the skull clamp adaptor and standard reference was applied securely. Registration was completed, achieving an accuracy of 0.5 mm, which was verified and accepted by the surgeon. During the procedure, real-time tracking of long probe was performed using optical guidance, ensuring precise targeting of the lesion. The navigation system responded accurately throughout the surgery, with stable reference and consistent tracking. Post-Operative Remarks: The system functioned smoothly with satisfactory accuracy. Surgeon appreciated the navigation performance and acknowledged successful support during the cranial biopsy procedure. | NA | NA | NA | NA | NA | Surgeons were satisfied and Impress with accuracy of system. | NA |  | NA | NA | Faiz Khan | Optical | V1.2.24 | Other |  | NA | 0.3 | 3 | 0.5 | https://drive.google.com/open?id=1c8zG0ulOe3aXWJsWAzZ4dQPSR8acOZM_ | https://drive.google.com/open?id=1JF4oVGgKmO91yJM4ifOo7Z3mRoGJPt8w | https://drive.google.com/open?id=1-xeVVVpKRsVeXKq7Q_MrMAqTmGWl4VUG
2025-10-29 21:31 | gurjit@rxoom.com | 2025-10-29 | Osmania General Hospital | Hyderabad | NA | Dr Nagarjuna, Dr. Laxman | Gurjit Singh | Cranial | Meningioma Tumor Excision |  |  | Yes |  | Completely | [5 empty] | NA | The pre-operative MRI scan was used. Skull tracker with mini plate was applied. Registration was performed successfully by surgeon, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA |  | Agree | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.22 | 5.49 | 0.32 | https://drive.google.com/open?id=1IY2hOcmMNCcQl6jhJZL_QCiQf73gRMwZ | https://drive.google.com/open?id=1b-X8BuLB1oXs6_gi4q56KDUzPM7cIVZc, https://drive.google.com/open?id=1JvjRHyl-Tbmozh0sU1kbKhYUDZT_02GJ | https://drive.google.com/open?id=18X3M0xGXAgUqUIQajdbCRLiF4FY05Olw
2025-10-31 20:30 | gurjit@rxoom.com | 2025-10-30 | Osmania General hospital | Hyderabad | NA | Dr. Srikant and Dr. Ullas | Gurjit Singh | CT-Spine | Spine fixation D9, D11, L2 |  | 145 | Yes |  | Completely |  |  |  | Orientation of patient scan was wrong. Patient anterior side was showing on posterior side in view and posterior side was on anterior side. |  | NA | Preoperative CT scan was loaded. In CT scan wrong orientation was identified, patient anterior side of spine was showing on posterior side of view and posterior side of spine was showing on anterior side, same was with left right. Understating the orientation issue, fiducial points were marked and capture patient on accordingly on single level vertebra. Single level vertebra registration was done for D9, D11 ,L2 before operating on them. Surgeon checked and confirm the trajectory and entry point of pedicle, confirm good accuracy. | Required other navigable spine tools in spine. |  | Strongly Agree | Strongly Agree | Strongly Agree | Easy Registration and upgrade to DBS navigation | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.08 | 3.07 | 0.24 | https://drive.google.com/open?id=1zn95wg41Wym3jv9shdl4OP2eojDE2Nfb | https://drive.google.com/open?id=1RnRqYxsYq0a0YveosFA9sgareInut8qw, https://drive.google.com/open?id=1N6GD3frfprdK67hOcWnAoyw1QL05V79b, https://drive.google.com/open?id=1dCBMuv8ukZbdt_Leh1NR70khwXCbgW9N | https://drive.google.com/open?id=15rcPhnu1M8OPVxkTIW5muVg1RB-bjDNS, https://drive.google.com/open?id=1_MAu96Nvff2bsWRoXeYSHz80E8WnN1W7, https://drive.google.com/open?id=1pwcvYTnLhOzFt2UORr6lEDO1k7cJdvAD |  |  | https://drive.google.com/open?id=1STNORkFw0Z6iG-c8ye4EPSEqqTHjUc_B
2025-10-31 21:22 | gurjit@rxoom.com | 2025-10-31 | Sindhu Hospital | Hyderabad | NA | Dr. Narasigna Rao, and Dr. Aneesh | Gurjit Singh | Cranial | Rt Occipital Meningioma |  |  | Yes |  | Completely |  |  | Software crash after working around the merging data, manually adjusting sliders to merge data, utilizing reregister option on merge data window and accepting the merged data. |  |  | NA | The pre-operative CT scan was used as the primary dataset, and the MRI scan was merged as secondary data. Patient position was prone and registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA |  | Good accuracy | NA | NA | NA | In tool tip calibration process tool tip should be calibrated only after a certain degree of tool rotation and angle of tool tip samples are given to system. Currently even tool is kept stable at same point gets tool tip get celebrated and only moving probe at some angle does not produce accurate tool tip calibration, after rotating tool we get the actual tool tip offset in probe. Other Tools like tracker tools should required tool tip calibration before using them like VP shunt tool with guidance needle and Biopsy guidance tool, upcoming tools. It will help ensure any changes in tool tip and factory tool calibration for testing, for user before using them or applying them in surgery. |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.09 | 4.16 | 0.51 | https://drive.google.com/open?id=1pLW3wEbCKNo_QrXjD1n-ZE8jvulOwWh5 | https://drive.google.com/open?id=11hAGRGOf_6M1phKDnYBOjCcLnmyhsgzP, https://drive.google.com/open?id=1f2TWmDktScNHz03teZrD3lBBxIWXuH7n, https://drive.google.com/open?id=1HJHrAEBtb1SXlFUMAkEZfBDsUIaJZvu_, https://drive.google.com/open?id=1OW8948GY4kbdbg8Z7tsPLyQf5psJYnlP | https://drive.google.com/open?id=1qJBkaMuhc2s-ChaU8pVYIyt3ZZqPGGYA
2025-11-01 21:45 | gurjit@rxoom.com | 2025-11-01 | Sindhu Hospital | Hyderabad |  | Dr. Narasinga Rao, Dr. Aneesh | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brain Biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and registration was done successfully, and accuracy was verified and accepted by the surgeon. Bur hole was made and surgeon aligned the biopsy guidance tool with pre planned trajectory and after applying the stopper on measured distance, surgeon took biopsies. Navigation was used completely throughout the surgery for guidance. | While trajectory planning surgeon view probe eye view of trajectory. While confirming through the trajectory on probe eye view surgeon suggested when the view/slices is gone beyond the target point the point of trajectory should not show. |  | NA | NA | NA | NA | NA |  | One glyonce coating | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 100 | 4.69 | 0.72 | https://drive.google.com/open?id=1HXZJesw66Q5Rsn_u1-A1ECogVLEIXjkN | https://drive.google.com/open?id=1SXYt-7nTzBbU8FJkeLpFKF8_HynXCnDr, https://drive.google.com/open?id=1goXRCK-1N1ZjklqOcC4r8hMlRKBQz224 | https://drive.google.com/open?id=1iRb1U_oKFMiwtVIrzFIqNPSDND0b5Jae
2025-11-03 21:19 | gurjit@rxoom.com | 2025-11-03 | Sindhu Hospital | Hyderabad | NA | Dr. Narasingha Rao, Dr. Aneesh | Gurjit Singh | CT-Spine | Spine fixation L2 and L4 |  | 10 | Yes |  | Completely |  | Universal geometry and spine clamp reference geometry causing reference recognition issue. Spine instrument working as reference and reference working as spine tool which results in sudden tool showing far in air while surgeon keeping the tool on spine. |  |  |  | NA | Pre operative CT spine was used and successful single level vertebrae registration was done on L2 and L4 by surgeon. New Spine instruments like awl, pedicle finder were used intra operative for precise pedicle screw placement. Navigation was used completely throughout the surgery for guidance. | Make small spine reference attachment. |  | Agree | Neutral | Strongly agree | NA | NA |  | Two Glyon coating came off | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.55 | 3.46 | 0.67 | https://drive.google.com/open?id=1vEp-ta5y_XSrNmZ65mk8lWiIBAm2dAIX | https://drive.google.com/open?id=1U08tybMwZ4gxfuv4_C0Avt8tOdUhgkTU, https://drive.google.com/open?id=1H0MvI_ozBYoYCIXrBrJy5Xnv-Z1pvtIP, https://drive.google.com/open?id=1XkYQ4PEQIbGV05u5J-arMdlbOWLPTGWy | https://drive.google.com/open?id=1Ys_RRi6lVmQlRZ6GuoXZ6KKmPdc4KRkq, https://drive.google.com/open?id=1WNHxE1jVtmsTToyF2zHaFDWaDkIuvTMX, https://drive.google.com/open?id=1jlQlxcBAS76twtpYy0ktSV9GKqajnZD7, https://drive.google.com/open?id=1AXcvae3Bzn1yYI0XSMgjIxDLGTcGegzR |  |  | https://drive.google.com/open?id=1OMXpM-95C_e-IfjIs_9IDztXOVqwIvHm
2025-11-03 21:50 | trushalgulhane1547@gmail.com | 2025-11-03 | Kiran hospital surat | Surat | NA | Dr. Ankit Patel | Trushal Gulhane and Krishna kumar | CT-Spine | L1 and L2 fixation |  |  | Yes |  | Completely | [5 empty] | NA | Date: 03 November 2025 Technology Used: Optical Navigation System Hospital: Kiral Hospital, Surat Surgeon: Dr. Ankit Patel Case: Spine Fixation – L1, L2 Accuracy Achieved: 0.92 mm Data Used: CT Procedure Summary: Pre-operative CT data was utilized for registration and surgical planning. The optical navigation system was set up successfully, and the CT clamp along with the CT reference was securely applied. Registration was completed with an achieved accuracy of 0.92 mm, which was verified and accepted by the surgeon. During the procedure, real-time tracking of the long probe, awl, tap, and pedicle finder was performed using optical guidance, ensuring precise alignment and targeting throughout the spine fixation process. The navigation system responded consistently and accurately. Fiducial Registration Accuracy: 6.24 mm Surface Registration Accuracy: 0.92 mm Duration of Procedure: 7:00 AM to 11:00 AM Post-Operative Remarks: The system performed efficiently with stable and satisfactory accuracy throughout the procedure. Dr. Ankit Patel appreciated the navigation system’s performance and acknowledged the successful technical support provided during the spine fixation surgery. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.2.24 | Other |  | NA | 0.2 | 6.24 | 0.92 | https://drive.google.com/open?id=18wq02JZoYlf1_jRAh8MXpMYy_LEG1jsV | https://drive.google.com/open?id=1noWtUx8ie4UNvqnpkDBez-NJt57hu9GZ | https://drive.google.com/open?id=11gBb8lRZpzZgDgTn-Ilak0TAeWXbLiCa
2025-11-06 22:51 | rohanghosh.2009@gmail.com | 2025-11-06 | Gouridevi Hospital | Durgapur | NA | Dr. Soumyadeep Nandi | Rohan Ghosh | CT-Spine | laminectomy | NA | 8 | Yes | NA | Completely | NA | NA | NA | NA | NA | Doctor found out our system works very good & handy | The NeuroSaphireX navigation system was successfully demonstrated during a laminectomy procedure at Gouridevi Hospital, Durgapur. Dr. Soumyadeep Nandi independently performed the registration and navigation process, achieving an accuracy of 1.02 mm. The surgeon expressed strong satisfaction with the system’s performance, specifically highlighting its accuracy, ergonomic design, and user-friendly interface. Dr. Nandi described the NeuroSaphireX as a very handy and useful system, noting that all functions operated smoothly throughout the surgery. No negative feedback was reported, and the overall response was highly positive, confirming the system’s reliability and ease of use in spinal navigation procedures. | Dr. Nandi described the NeuroSaphireX as a very handy and useful system, noting that all functions operated smoothly throughout the surgery. No negative feedback was reported, and the overall response was highly positive, confirming the system’s reliability and ease of use in spinal navigation procedures. | NA | System of the accuracy is the best part of the system | Strongly agree | Strongly agree | Very handy & very useful instrument | NA |  | NA | NA | Rohan Ghosh | Optical | 1.2.24 | NS TRL 03 SE 101 | Na | Na | 0.14 | 4.96 | 1.22 | https://drive.google.com/open?id=1GUgHp_lPbI6yVe5IkDFhiprPix396fiU | https://drive.google.com/open?id=1yD8oyb68XRBfTfTn69z9VZa1VJEblgiX, https://drive.google.com/open?id=1yB5Sa_S3qqvNGJDuG1TwErM3XAEpTJNC, https://drive.google.com/open?id=1LL7jOKSVAw1v7k9OmZlmVUKU_J-BzNZr | https://drive.google.com/open?id=1tjFBo7T99mRu5f-ImNpAoEZM6rnOHh53 | https://drive.google.com/open?id=19_4b4JwDprlPem_Ncro0nX79x1hFNmkj |  | https://drive.google.com/open?id=1jzH7Hmcmy4K_T388dLLUH1xEkvmLwadj
2025-11-07 19:21 | trushalgulhane1547@gmail.com | 2025-11-07 | Kiran hospital | Surat | Na | Dr. Bhaumik thakor and dr. Ankita patel | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | Na | Hospital: Kiran Hospital, Surat Date: 7th November 2025 Procedure: Brain Biopsy Surgeons: Dr. Ankit Patel, Dr. Bhaumik Thakor Technology Used: Optical Navigation Technology Procedure time : 2.30 hours Pre-operative Details Patient MRI data was provided by the OT staff. Initially, the CD drive was unable to read the data. Upon request, the OT staff provided a compatible and acceptable data format. The MRI data was successfully loaded into the system, and an accurate 3D model was generated for planning purposes. Planning Trajectory planning for the brain biopsy was performed by Dr. Bhaumik Thakor and Dr. Ankit Patel. The following accuracy parameters were achieved: Fiducial Registration Accuracy: 4.45 mm Surface Registration Accuracy: 0.98 mm Instrument setup and system registration were completed successfully with an overall accuracy of 0.9 mm. Intra-operative Details After sterile draping, tooltip calibration was performed. The biopsy tracker guide was aligned precisely with the planned trajectory. A projection was displayed along the biopsy tracker guide up to the target point. The projection length and biopsy tracker length were calculated. This measurement was marked on the biopsy needle, and a stopper was placed accordingly. During the insertion, it was observed that the biopsy needle extended slightly beyond the planned target. To verify real-time tracking, the system was switched to the biopsy needle tracker, which accurately displayed live tracking of the needle. The biopsy was then completed successfully. Summary Patient data successfully loaded and 3D model generated. Planning and registration performed with acceptable accuracy. Live tracking confirmed accurate needle positioning. Procedure completed successfully using optical navigation technology. | Na |  | Na | Na | Na | Na | Na |  | Na | Na | Faiz Khan | Optical | V1.2.24 | Other |  | Na | 0.3 | 4.45 | 0.98 | https://drive.google.com/open?id=11paR59x-54-EyWOgRJWqkYGZzbeI1w3D | https://drive.google.com/open?id=1yw-9vjVUYYm59R7No5JBgdgIGnp1UhLV | https://drive.google.com/open?id=19Qltu8IDyB6pUSuYCVThScPp71hU5hVZ
2025-11-07 22:33 | rohan@rxoom.com | 2025-11-07 | Gouri Devi Hospital | Durgapur | NA | Dr. Saikat Saha | Rohan Ghosh | CT-Spine | Dorsal Spine Fusion | Dorsal Spine Fusion | 2 | Yes | NA | Completely | NA | NA | NA | NA | NA | Dr. Saikat Saha successfully utilized the NeuroSaphireX Navigation System during a dorsal spine fusion procedure. The doctor appreciated the system’s ease of use, mentioning that it is “easy to handle and use.” He also found the navigation process accurate, the user interface simple, and the instruments ergonomically designed, which contributed to a smooth surgical experience. In his overall feedback, Dr. Saha described the system as “very helpful”, emphasizing its clinical utility and operational simplicity. The positive response reflects strong surgeon confidence in NeuroSaphireX’s accuracy, efficiency, and user-friendly design, making it a valuable tool for spine navigation surgeries. | The NeuroSaphireX navigation system was successfully utilized during a Dorsal Spine Fusion procedure performed by Dr. Saikat Saha. CT-based navigation was employed for this case, ensuring precise localization and guidance throughout the surgical process. The overall navigation workflow—from patient data import to accuracy validation—was executed smoothly. Dr. Saha appreciated the ease of handling, simple registration process, and user-friendly interface. The system’s accuracy and ergonomic instrument design were specifically acknowledged as major strengths. The surgeon described the system as “easy to handle and use” and provided an overall positive impression, noting that the system was “very helpful” for spine procedures. Observations & Feedback: The navigation accuracy remained stable throughout the case. Surgeon noted that sometimes the system did not detect the weapon/instrument, which could be due to instrument tracking sensitivity or reflective marker alignment issues. Despite this, the overall workflow was satisfactory and supportive for clinical use. Suggestions for Improvement: Enhance tracking sensitivity: Improve camera calibration and marker reflection detection, especially under high illumination or variable angles. Instrument auto-detection alert: Introduce a prompt when an instrument is not detected or temporarily out of range. Workflow Optimization: Provide a brief on-site system check protocol before surgery to verify instrument visibility and camera alignment. Training Support: Continue to educate clinical users on maintaining marker cleanliness and optimal setup angles for consistent detection. Conclusion: The demonstration at Gouridevi Hospital was successful, and Dr. Saikat Saha expressed high satisfaction with the NeuroSaphireX system’s performance. The system’s accuracy, user-friendliness, and ergonomic design were key highlights. With minor refinements in tracking reliability, the system can further enhance its consistency and strengthen its position in the competitive neuronavigation market. | NA | NA | Strongly agree | Strongly agree | Strongly agree | Over ally very helpful | NA |  | NA | Easy to handle & use | Rohan Ghosh | Optical | 1.2.24 | NS TRL 03 SE 101 | NA | NA | 0.15 | 4.38 | 0.84 | https://drive.google.com/open?id=1V1tvKe693e0-BBhHIR34rgLptXT1BV9F | https://drive.google.com/open?id=1UDspJJvy8Kmt1iSmjoNxbHP3AqN2O46h, https://drive.google.com/open?id=1ZNN4c1C1qeEdqhdS5qqRMLqWdlWMJt8w, https://drive.google.com/open?id=1ovi-C-bvRUdTzOGQQfhJn4xiIg5uRS3J | https://drive.google.com/open?id=1ZUP7ZWsGA8KhPfE0fReoCv74rlSNZU65 | https://drive.google.com/open?id=1WTB1JeDTsjD3u8AvOop2dH6BKFYRWk3c |  | https://drive.google.com/open?id=1w1tPl90PznrGLTddrErAQl7Cf-MGZRd8
2025-11-08 22:55 | rohan@rxoom.com | 2025-11-08 | Gouri Devi Hospital | Durgapur | NA | Dr. Soumyadeep Nandi | Rohan Ghosh | Cranial | Craniotomy | Frontal Craniotomy | 0 | Yes | NA | Completely | NA | NA | NA | NA | NA | Dr. Soumyadeep Nandi successfully used the NeuroSaphireX Navigation System during the craniotomy procedure. He appreciated the ease of usage, user-friendly interface, and simple workflow of the system. The data loading and registration process were noted as smooth and efficient. However, the surgeon observed a slight hold in accuracy even after achieving satisfactory registration, suggesting a potential inconsistency between registration and intraoperative navigation accuracy. | Thank you for providing the doctor's feedback form. This document contains very useful information regarding the use of the NeuroSaphireX (Neuro Navigation System) during a surgery at Gouri Devi Hospital in Durgapur. Here is a summary and analysis of the feedback, including significant points for system improvement and demo practices. Summary Report on NeuroSaphireX System Use Key Surgery and System Details ¦ Aspect ¦ Detail ¦ ¦ **Hospital Name & Location** ¦ Gouri Devi Hospital, Durgapur ¦ ¦ **Date of Surgery** ¦ 08/11/2025 ¦ ¦ **Surgeon Name** ¦ Dr. Soumyadeep Nandi ¦ ¦ **Surgery Details** ¦ Craniotomy ¦ ¦ **RXoom Representative** ¦ Mr./Ms. Rohan Ghosh ¦ ¦ **Surface Registration Accuracy** ¦ $0.35\text{ mm}$ ¦ System Performance Feedback The feedback indicates a highly positive experience with the NeuroSaphireX system, with the doctor **Strongly Agreeing** with four out of the five key statements: ¦ No. ¦ Aspect/Statement ¦ Agreement Level ¦ ¦ 1 ¦ It was easy to load patient data ¦ **Strongly Agree** ¦ ¦ 2 ¦ The registration process was simple ¦ **Strongly Agree** ¦ ¦ 3 ¦ Navigation was accurate ¦ **Agree** ¦ ¦ 4 ¦ Instruments were ergonomic ¦ **Strongly Agree** ¦ ¦ 5 ¦ The User interface (UI) was simple ¦ **Strongly Agree** ¦Analysis and Significant Points This successful demo and high score provide an excellent testimonial. The following points are significant for both product improvement and optimizing demo/sales practices: ### I. High-Value Selling Points (What worked well) * **Exceptional Ease of Use:** The surgeon rated four out of five factors as "Strongly Agree," specifically noting that the **"Ease of usage"** was the feature/aspect of the device they liked the most. This is a crucial selling point for any complex medical device. * **Simple Setup and Workflow:** Statements 1, 2, and 5 (Easy data load, simple registration, simple UI) all reinforce that the workflow is efficient and non-disruptive, which is highly valued in a fast-paced operating room environment. * **High Registration Accuracy:** The recorded **Surface Registration Accuracy of $0.35\text{ mm}$** is an extremely impressive technical metric that should be highlighted prominently in future marketing and demo materials. This is a quantifiable measure of precision. * **Ergonomics:** The "Strongly Agree" for ergonomic instruments indicates that the physical design is well-suited for surgical procedures. ### II. Areas for System/Demo Improvement (The 'Agree' point) * **Navigation Accuracy:** The only statement that received a rating of **"Agree"** instead of "Strongly Agree" was **"Navigation was accurate."** * **Actionable Point:** While "Agree" is still positive, this suggests there might be a subtle perception that the navigation could be *more* accurate or that the surgeon may have encountered a minor, non-critical deviation. This should be investigated: * **System Check:** Review the system log for this specific case for any potential errors or drift during the procedure. * **Demo Practice:** The representative (Rohan Ghosh) should follow up with Dr. Nandi to understand what would have elevated this rating to "Strongly Agree." This qualitative feedback is vital. ### III. Surgeon's Written Feedback The written comments offer further, targeted insights: * **Liked Feature:** "Ease of usage." (Confirms the strongest positive aspect). * **Feature/Aspect *Not* Found Useful:** **"Mild inaccuracy on side of satisfactory registration."** * **Actionable Point:** This written comment directly relates to the "Navigation was accurate" rating. It suggests that while the **initial registration was satisfactory/simple**, there was a subsequent **"mild inaccuracy"** on the side (periphery) of the registered area. * This is the **most significant technical point for system improvement.** The R&D team should focus on ensuring uniform and robust accuracy across the entire operative field, especially at the margins/periphery, after the initial satisfactory registration. ### ✅ Next Steps and Recommendations 1. **Follow-Up on Accuracy:** Have the technical team analyze the system logs for the $0.35\text{ mm}$ registration accuracy case to understand the cause of the reported "mild inaccuracy on side" (Dr. Nandi's concern). 2. **Highlight Key Metrics:** Use the $0.35\text{ mm}$ accuracy and the "Strongly Agree" scores for **"Ease of Usage," "Simple UI,"** and **"Ergonomics"** as primary case study evidence for future sales and marketing efforts. 3. **Ensure Promotional Consent:** Since the surgeon has checked the consent box to use the feedback for promotional purposes, this document (and the positive scores) can be immediately leveraged for testimonials. | NA | NA | Strongly agree | Strongly agree | Strongly agree | NA | NA |  | NA | NA | Rohan Ghosh | Optical | 1.2.24 | NS TRL 03 SE 101 | NA | NA | 0.07 | 3.71 | 0.93 | https://drive.google.com/open?id=1VRRZV3h99fcotOLtGBTvBG_DyUc2C6Nt | https://drive.google.com/open?id=15bcn2CaqP0mZns4SddNgqhnpAnm-JePh, https://drive.google.com/open?id=1H-onunDMh5X84pdA8gc6Pt6S6Lqbkvlk, https://drive.google.com/open?id=1O1ahBP9t4ZY_4USpBUfmJtR85Oz0RicC | https://drive.google.com/open?id=1sjerKfwV0OxFMXfoITOIcQktaJJhD2ib | https://drive.google.com/open?id=1V8nAO8UVDebC2qaEb8xxOjNgwSuopHS0 |  | https://drive.google.com/open?id=1bg3vzEGVAVHES2dLGPuC-Vnei6dgj4_i
2025-11-10 21:59 | trushalgulhane1547@gmail.com | 2025-11-10 | Kiran hospital | Surat | NA | Dr. Bhaumik Thakore and dr. Rahul Patel | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | Na | Hospital: Kiran Hospital, Surat Date: 10th November 2025 Procedure: Endoscopic TNTS Surgery ENT Surgeon: Dr. Rahul Patel Neurosurgeon (HOD): Dr. Bhaumik Thakor Technology Used: Optical Navigation System Approximate Procedure Duration: 3 hours 30 minutes Pre-Operative Details Patient CT scan data was provided by the OT staff and successfully loaded into the navigation system. A precise 3D anatomical model was generated for preoperative planning and registration purposes. For this case, a three-pin head fixation was not used. To maintain navigation accuracy during patient movement, it was recommended to use a skull reference tracker. However, since the surgeon preferred not to screw the tracker to the patient’s skull, the tracker was instead secured using Dynaplast to ensure adequate stability and minimize movement. Registration was then performed successfully with the following results: Fiducial Registration Accuracy: 5.08 mm Surface Registration Accuracy: 1.3 mm The registration accuracy was verified and found acceptable by the surgeons. Intra-Operative Details Following sterile draping, the unsterile reference tracker was replaced with a sterile one, and tooltip calibration was completed successfully. The system maintained consistent accuracy and stability throughout the procedure. Post-Operative Remarks After the successful completion of the surgery, Dr. Bhaumik Thakor expressed appreciation for the accuracy and stability of the navigation system during the procedure. Summary Patient CT data was successfully imported and processed to generate a 3D model. Registration was performed with clinically acceptable accuracy. Real-time tracking was stable and precise throughout the surgery. Procedure completed successfully using Optical Navigation Technology. | Na |  | Good accuracy | Na | Na | Na | Na |  | Na | Na | Faiz Khan | Optical | V1.2.24 | Other | W | Na | 0.2 | 5.08 | 1.33 | https://drive.google.com/open?id=1cgSFzsW3i3LQkJ14lDSkrgwuMAJPDjaE | https://drive.google.com/open?id=1eLCnBBeupyqejtWD7MeAsIZC_twjvsLg | https://drive.google.com/open?id=1z16T13F2bxbFZPHGV9GKTwVrv-i11n3p
2025-11-11 13:32 | trushalgulhane1547@gmail.com | 2025-11-11 | Kiran hospital | Surat |  | Dr Ankit patel | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | Na | Hospital: Kiran Hospital, Surat Date: 11th November 2025 Procedure: VP Shunt Surgery Neurosurgeon: Dr. Ankit Patel Technology Used: Optical Navigation System Approximate Procedure Duration: 2 hours 30 minutes Pre-Operative Details Patient CT scan data was provided by the OT staff and successfully loaded into the Optical Navigation System. A detailed 3D model was generated for preoperative planning and registration purposes. In this case, a three-pin head fixation was not used. To maintain navigation accuracy during patient movement, the use of a skull reference tracker was recommended. However, as the surgeon preferred not to screw the tracker to the patient’s skull, the tracker was securely fixed using Dynaplast to ensure adequate stability and minimal movement. Registration was then completed successfully with the following accuracy values: Fiducial Registration Accuracy: 3.21 mm Surface Registration Accuracy: 0.62 mm The registration accuracy was verified and found satisfactory by the operating surgeon. Intra-Operative Details After sterile draping, the unsterile reference tracker was replaced with a sterile one, and tooltip calibration was completed successfully. The VP shunt guide was used for navigation to the targeted lesion site, followed by precise insertion of the shunt tube. The system maintained consistent accuracy and stability throughout the procedure. Post-Operative Remarks Upon successful completion of the surgery, Dr. Ankit Patel appreciated the navigation accuracy and expressed satisfaction with the system’s performance. Summary Patient CT data was successfully imported and processed to generate an accurate 3D model. Registration was performed with clinically acceptable accuracy. Real-time navigation and tracking were stable and precise throughout the procedure. Surgery completed successfully using Optical Navigation Technology. | Na |  | Good feedback | Good | Good | NA | NA |  | NA | Na | Faiz Khan | Optical | V1.2.24 | Other |  | NA | 0.2 | 3.21 | 0.62 | https://drive.google.com/open?id=1ff2aLk0gDH5qVcgY3V-s5EoN2Y-iCWl4 | https://drive.google.com/open?id=1uv84hYs5ZpPBlkiycJRxdfIp-3Sk8dEC | https://drive.google.com/open?id=1XiITQpwDR6Lt0P3SccfYbsms7RvLWXs1
2025-11-14 21:25 | gurjit@rxoom.com | 2025-11-14 | TX Hospital | Hyderabad |  | Dr. Naresh Kumar, Dr. Chandra Shekav | Gurjit Singh | Cranial | Craniotomy Tumor Excision |  |  | Yes |  | Completely |  |  |  |  | Their 3pin head clamp was made in Indian, we initially try to attach reference arm to it but their clamp had less thread than the standard clamp, and our reference arm knob was not getting fully tight so we used the table adapter. | NA | Preoperative CT scan was loaded. Patient position was prone and registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | NA | Reduce threads on knob of standard reference arm, so that its direct attachment can increase on different types of 3 pin head clamp. |  | one glyonce coating | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.14 | 7.05 | 0.84 | https://drive.google.com/open?id=1iGxzzYjDjm7bfQmMi807-wjk7cIJIg5C | https://drive.google.com/open?id=12BakhgIANIvYMiduvjWN2tvu6bJ7z_SO | https://drive.google.com/open?id=1zlllp6aLqL8OkuSDLVBaJS19a_kInlF2 |  |  | https://drive.google.com/open?id=16Ta9CRxwTy_Qoq5-tPaSE40R8yQRu6lQ
2025-11-20 21:11 | trushalgulhane1547@gmail.com | 2025-11-18 | Mahatma Gandhi hospital | Jaipur |  | Dr. Pankaj Gupta | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | During the procedure navigation accuracy affected by slightly movement of skull reference |  |  |  |  | NA | Post-Operative Remarks Post-surgery, a discussion was held with the operating team. It was acknowledged that the accuracy deviation occurred due to movement of the skull reference tracker. It was conveyed that fixing the tracker with a screw would significantly minimize such issues and enhance accuracy. The surgeon agreed with this but mentioned that it is not feasible for them to use a skull-screwed tracker in every case. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.3 | 5.53 | 1.12 | https://drive.google.com/open?id=1AzWIQXubbQv12cqCf46_mfN5hqM2JKM3 | https://drive.google.com/open?id=1-fiFcD1UWX8ZD0-Qok5zCmmnEeojiUu7 | https://drive.google.com/open?id=1Fdvwg-eOzshXWhJflE4iXY6oVCeeg19x
2025-11-20 22:21 | trushalgulhane1547@gmail.com | 2025-11-19 | Mahatma Gandhi hospital | Jaipur |  | B.S. Sharma | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | System used for incision planning |  |  |  |  | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.4 | 2.22 | 0.43 | https://drive.google.com/open?id=1jNSLLk_F_krFvT2cbOKyxUi7D3Oygohk | https://drive.google.com/open?id=1letCpYJx8zXRoJiA6mZXpEI8LoCHIoPp | https://drive.google.com/open?id=1JJvP1Y3crT_YXSMKFrkg25OtphE-f--f
2025-11-20 22:27 | trushalgulhane1547@gmail.com | 2025-11-20 | Mahatma Gandhi hospital | Jaipur |  | Dr. Jitendra | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successfully used the system | Mini location block design in such away must be covered by skull reference |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 10.8 | 1 | https://drive.google.com/open?id=1_VFtEtbVU37rrjIDHGzn7e50sUVDLNpa | https://drive.google.com/open?id=1p9qOKXyBFYUb0O_L48w3eNtmGyv_u9gF | https://drive.google.com/open?id=1zh3vzt7SdExXV5K5zwldYX7bJgq0YAqM
2025-11-21 19:40 | trushalgulhane1547@gmail.com | 2025-11-21 | Mahatma Gandhi hospital | Jaipur |  | Dr. Bhatia | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successfully used the system | NA |  | Good | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V2.2.24 | Other |  | NA | 0.2 | 3.6 | 0.69 | https://drive.google.com/open?id=1vZcEB1lJkviJ6cDJbidSZAevrX10yAy5 | https://drive.google.com/open?id=1OqMn4TqpfftW2cEtQ97R_O3Hvo5Gb_K6 | https://drive.google.com/open?id=18lODzvQNZy7u1dLrLqYbY8vJoobKfS4F
2025-11-26 17:39 | trushalgulhane1547@gmail.com | 2025-11-26 | Mahatma Gandhi hospital | Jaipur |  | Dr. B.s. pankaj Gupta | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.4 | 3 | 0.8 | https://drive.google.com/open?id=1GXZ_1pEHwBpMs59gUScdXLlHJPLfxTCy | https://drive.google.com/open?id=1hpbxZgv9E6ALEJoLJK_UbgrD80nBP_hf | https://drive.google.com/open?id=1nV1dUC6slAHcGZKp2eYx_l1yQJ3XQe5h
2025-11-27 17:55 | trushalgulhane1547@gmail.com | 2025-11-27 | Mahatma Gandhi hospital | Jaipur | NA | Dr. Jitendra | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | [5 empty] | NA | System used for incision planning | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.4 | 4.1 | 1.28 | https://drive.google.com/open?id=1rH9kHf2gJu6piAI5cSbc4z-Tz50h2-KE | https://drive.google.com/open?id=1v8xhCXVy5xRrrXFxGTWMmc8z9VeDpyc1 | https://drive.google.com/open?id=19gXU1JufBz_kcF_9t_hsIqDkkrS2Lhlk
2025-11-27 18:10 | trushalgulhane1547@gmail.com | 2025-11-27 | Mahatma Gandhi hospital | Jaipur |  | Dr. Bhatia | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | System successfully used by dr. Bhatia | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 1.58 | 0.39 | https://drive.google.com/open?id=1FO3P43m38pdAgIFq64_LOUWVCWgaopzn | https://drive.google.com/open?id=1kK8IDbDqXHeFBOBu91dzlxvyqUc7gEy9 | https://drive.google.com/open?id=1GBeAag0AGvl1kM3G6E2t5QG7G0A4AD3s
2025-11-29 22:41 | trushalgulhane1547@gmail.com | 2025-11-29 | Mahatma Gandhi hospital | Jaipur | NA | Dr. B.s. sharma and dr.jitendra | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successfully used the neuro navigation system | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 2.59 | 0.45 | https://drive.google.com/open?id=17bys-6W75V4zI-4rX4eSL9FQTWg6OmBr | https://drive.google.com/open?id=1uGab0Ljh8yCxALG5JeteiGodoTYcb0Ja | https://drive.google.com/open?id=1ELmC0IhYyWG5lU4h9KTqxD9C8xVKKxq7
2025-12-04 20:44 | gurjit@rxoom.com | 2025-12-04 | JNMC | Aligarh | NA | Dr. Raman Mohan Sharma, Dr. Ahmed Ansari | Gurjit Singh | CT-Spine | L3-L5 Spine fixation |  | 30 | Yes |  | Partially |  | NA | NA | NA | During system installation it was going after the train transportation of trolley from Hyderabad to Delhi the trolley lower pole top is damaged which cause issue in installation. Damage was also was also on trolley camera arm mounting pole bolt top threads which cause nut got stuck in there. | NA | Preoperative CT scan was loaded. Surgeon decided the fiducial points and points were marked and capture patient on accordingly on L3 single level vertebra. Surgeon checked and confirm the trajectory and entry point of pedicle, confirm good accuracy one side of pedicle and slightly inaccurate on other side of pedicle. | NA |  | Agree | Strongly Agree | Agree | More Trial should be conducted at Govt hospital to make is more user accessible and friendly. | NA |  | Two Glyonce coating | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.13 | 2.56 | 0.25 | https://drive.google.com/open?id=1YlujZ7U3j2mXWeFSEypbTq3iGk6o5qPj | https://drive.google.com/open?id=1b_pXDRLsPs6-9Ov2OaLJj22rONX0y3Xo, https://drive.google.com/open?id=1L5lHp5sSo0HMBAaqU0LzFziFXT2NGy-F, https://drive.google.com/open?id=1yahRUe02_rd2tEy-lr1mGlnHWAVTgSKc, https://drive.google.com/open?id=1wP6xawfUgIqOhetpuyKiUpQM8itKK6d9 | https://drive.google.com/open?id=1RqXMc69OxB8mIFv0T9l8qndcG12zYpoZ, https://drive.google.com/open?id=1jMXJjR6-fqieEcSavY8BCCcDBK3HVeu9 |  |  | https://drive.google.com/open?id=12QLsHao81jVySD-jp_sAEhI6dIbQ-akY
2025-12-08 06:15 | trushalgulhane1547@gmail.com | 2025-12-07 | N.c.poonia neuro care | Jaipur |  | Dr.N.C.poonis | Trushal Gulhane | Cranial |  |  |  | No |  | Partially | [5 empty] | NA | After incision there was accuracy deviation | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.1 | 1.58 | 0.63 | https://drive.google.com/open?id=1uLqLEz94Eyk6wcAaFucK6GIOeog9_vMQ | https://drive.google.com/open?id=1IFtcLpDKiRj6RK9P3TqoBHsAmBw2y9W3 | https://drive.google.com/open?id=1VHGyTzp1XQX4V1f1iEJ6y5HT23Uf3pp6
2025-12-09 21:32 | trushalgulhane1547@gmail.com | 2025-12-09 | N.c.poonia neuro hospital | Jaipur |  | Dr. N.C. poonia | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.1 | 3.04 | 0.54 | https://drive.google.com/open?id=1cOoHmD5BCeLPY1vWj7LA9PUAW6xUBloJ | https://drive.google.com/open?id=1z6q4mboBNu_l6zCrwf-RpRjH0vz4td87 | https://drive.google.com/open?id=1-LTomGOZID1gj2NWNYTL6Dlrt64lX9s1
2025-12-29 11:44 | trushalgulhane1547@gmail.com | 2025-12-23 | Ramchandra Hospital | Dharmpur |  | Dr. Shreyansh choudhary | Trushal Gulhane | CT-Spine |  |  |  | Yes |  | Partially | [5 empty] | NA | AP view was good but face issues in other view | NA |  | Ok need multiple registration process | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 12.56 | 1.03 | https://drive.google.com/open?id=1iPR0ZDuX7G4PDF7eDzN9XnVzFoQZAQF9 | https://drive.google.com/open?id=1aCIabVzqUb9y2rBH_nOyBw2Xg47BGJRI | https://drive.google.com/open?id=1wysjdNVb05U5JLRBfXWEAT-L5E45Y_gT
2025-12-29 11:54 | trushalgulhane1547@gmail.com | 2025-12-24 | Ramchandra Hospital | Dharmpur | NA | Dr. Shreyansh choudhary | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.12 | 4.25 | 0.45 | https://drive.google.com/open?id=1Z2s2rMegnWR7NYNO-CUx42-HGJ1tLEG3 | https://drive.google.com/open?id=1Ifhezh9QH3CL8FCTMw9ZqDCB6Ll5s7AX | https://drive.google.com/open?id=1tKjsi_FIw7RdeQXNphrEMzcPf4FHWjQh
2025-12-29 15:11 | trushalgulhane1547@gmail.com | 2025-12-26 | Ramchandra Hospital | Dharmpur |  | Dr. Shreyansh choudhary | Trushal Gulhane | CT-Spine |  |  |  | Yes |  | Partially | [5 empty] | Length of spine clamp need to be reduced | Surgeon stated that in AP view was good but in later there was error deviation observed observed | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.1 | 8 | 0.4 | https://drive.google.com/open?id=1B674XYUv04y6zCZuMeJTqFeAHEQ5F25E | https://drive.google.com/open?id=1H_nzIBW8d273kG1LAxg6-sjpZ7KNgXK0 | https://drive.google.com/open?id=1VZwxZuFG-8IeUi8xmwiM7Pfwu-kOoWAK
2025-12-29 15:27 | trushalgulhane1547@gmail.com | 2025-12-29 | Ramchandra Hospital | Dharmpur |  | Dr. Shreyansh choudhary | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | Successfully used | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.3 | 4.02 | 0.45 | https://drive.google.com/open?id=1I91wD56GUYJQ_b-mMNvWf00UDmH2vY4V | https://drive.google.com/open?id=1g-zx-vJa6pPeAgqRo_HZZijEeAG8aiIh | https://drive.google.com/open?id=1LSvb3MYQOiNrFx-YOL-328X1wmOqtgYt
Demo Details 2026
2026-01-22 22:32 | gurjit@rxoom.com | 2026-01-22 | Unique Hospital | Surat |  | Dr Harin modi | Gurjit Singh | X-Ray Spine | Spine Fixation L4-5 |  | 1000 | Yes | In X-ray based spine surgery after incision spine was reregistered still accuracy was off 1 to 2cm according to the surgeon. | Not used | In X-ray based spine surgery after incision spine was reregistered still accuracy was off 1 to 2cm according to the surgeon. |  |  |  |  | Before incision ap and lat X-ray shot of patient were taken and loaded into the system and registration was done the. Spine was opened and then for re registration ap and lateral was taken again and re registration was done again. Then Dr checked accuracy and reported 1 to 2 cm error and compared tool Position in X-ray vs showing on navigation. Dr also reported last time same 1 to 2cm error was faced. | NA | NA |  | Accuracy if off by 1 to 2 cm | NA | NA | In previous demo system error was same 1 to 2 cm | NA |  | NA | Na | Trushal Gulhane | Optical | 1.2.24 | Other | NS TRL 04 DB 101 | NA | 100 | 100 | 0.41 | https://drive.google.com/open?id=1Rf6OIFhCRQGDQxi-4uEDLiIwSnqZ03e4 | https://drive.google.com/open?id=1OM7lP7oWlnzldDgp8l9kya9AoR9Yqjgk | https://drive.google.com/open?id=1cK1lHtI1S0q83I28Q2lEDefEI9LaOdTy
2026-01-23 16:30 | gurjit@rxoom.com | 2026-01-23 | Aaihms Hospital | Surat | NA | Dr Maulik Patel | Gurjit Singh | Cranial | Tumor Removal |  | 1000 | Yes |  | Completely | [5 empty] | Preoperative MRI scan was loaded, skull tracker with mini plate was applied. Patient registration was performed successfully by surgeon, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA | NA |  | Strongly agree | Strongly agree | Strongly Agree | Liked user interface and skull tracker | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | Other | NS TRL 04 DB 101 | NA | 0.17 | 3.41 | 0.61 | https://drive.google.com/open?id=1boe0l2U9SenvjJlum7FQbX44qiLLoChx | https://drive.google.com/open?id=1--z7GGBZv5L3PYLGV8JNnhr61ir9_aiy, https://drive.google.com/open?id=1U5WYtJFtzUYng0aE37D20POu-pdbyJ17 | https://drive.google.com/open?id=1tcRU6uxk4Tx3Ov9N_Tx0PQrYo6mxLzHe | https://drive.google.com/open?id=1lKO1t8GUPDKaQCF0VqWB48CYN8nkcDGE |  | https://drive.google.com/open?id=1KbR0irz74fDuNFbG1vOv38SZhjKGsQMR
2026-01-29 21:16 | gurjit@rxoom.com | 2026-01-29 | Aaihms Hospital | Surat | NA | Dr. Hitesh Chitroda | Gurjit Singh | Cranial | Tumor Excision |  |  | Yes |  | Completely | [5 empty] | Preoperative MRI scan was loaded, skull tracker with mini plate was applied. Patient registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | Registration process was easy and more accurate | Have diagonally pasted titanium mini plate under the base block instead of linearly. |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | Other | NS TRL 04 DB 101 | NA | 0.12 | 3.98 | 0.61 | https://drive.google.com/open?id=1Vm6Hq-zNPKSOcB5cxVvOnfXd5dmL7EUg | https://drive.google.com/open?id=1YBfty4UVBN5Htm6SPaVnFmdEBNWk9rrI, https://drive.google.com/open?id=17kncYzs1tu9xvkto1KxUKTjAgr-IRQz1 | https://drive.google.com/open?id=107Lq2XN7AoEJVD2n_gzeIh0zJAXBkcaC, https://drive.google.com/open?id=13wwH85J0evNJ7jkPPqq1SVRP16GRSJPS, https://drive.google.com/open?id=1bsIkvIqzR_HVWtZ2DSa6kg3m42Ov6lnB |  |  | https://drive.google.com/open?id=18zgQKNe6jiP9DGHIUGIVbU0T98_99vx9
2026-01-31 21:19 | gurjit@rxoom.com | 2026-01-31 | Aaihms Hospital | Surat | NA | Dr. Hitesh Chitroda | Gurjit Singh | Cranial | Tumor Excision |  |  | Yes |  | Completely | [5 empty] | Preoperative MRI scan was loaded, skull tracker with mini plate was applied. Patient registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation. | NA | Make more smaller prober than existing probe to use it while using microscope. |  | Strongly Agree | Strongly Agree | Strongly Agree | Make more smaller prober than existing probe to use it while using microscope. | NA |  | 4 Titanium Screw | NA | Gurjit Singh | Optical | 1.2.24 | Other | NS TRL 04 DB 101 | NA | 0.37 | 3.78 | 0.4 | https://drive.google.com/open?id=140OY5Dx8m-k8_XdfvYxncr05MdK03wvY | https://drive.google.com/open?id=1uMojUGEijHfffwdWwFEWzpdHPOSKFIt6, https://drive.google.com/open?id=1viu_F5cxElzEL7OLif1cRfa6g9ZD2CqX | https://drive.google.com/open?id=1-FmUPSyoS0tW0G85Hj_rbWVoOvP84i1D |  |  | https://drive.google.com/open?id=14sIRqp-711kb2Kwm99sdPc-iltYNlTPl
2026-03-17 10:17 | trushalgulhane1547@gmail.com | 2026-01-01 | Unique hospital surat | Surat | NA | Dr. Haren Modi | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Surgical Report Date: 1 January 2026 ▼ Hospital: Unique hospital surat, Gujarat Procedure: Aneurysm clipping Operating Surgeon: Dr. Haren Modi Imaging Modality: mri and flair data Surgical Outcome: Successful Navigation Accuracy: 1.11mm Case Summary: procedure conducted at Unique Hospital, surat, Gujarat, we began by utilizing mri and flair data provided by the OT staff. This data was imported into the system, generate a 3D model at the planning station. The surgical planning was then completed based on this model and data. after planning, instrument setup was carried out, the registration process was performed by surgeon, and. The system maintained a stable accuracy of | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 3.5 | 1.11 | https://drive.google.com/open?id=1zpxk2cZgfsJvUqLcGEfxY-s6YYFOIgli | https://drive.google.com/open?id=1S_AuhatgMMNiab1EyZKPep0m9iGBtB2s | https://drive.google.com/open?id=1HQ9ojn-7DD3Rb-gJ07l0PPtRouSUC0f9
2026-03-17 10:21 | trushalgulhane1547@gmail.com | 2026-01-06 | BAPS Hospital | Surat | NA | Dr. Haren Modi | Trushal Gulhane | CT-Spine |  |  |  | Yes |  | Completely | [5 empty] | NA | Surgical Report Date: 6 January 2026 Hospital: BAPS Hospital, Surat, Gujarat ▼ Procedure: Cervical Spine Fixation Operating Surgeon: Dr. Haren Modu Imaging Modality: CT Data Surgical Outcome: Successful Navigation Accuracy: 1.44 mm Case Summary: The procedure was conducted at BAPS Hospital, Surat, Gujarat. Preoperative CT data provided by the OT staff was utilized for surgical planning. The CT data was successfully imported into the neuro navigation system, and a three-dimensional anatomical model was generated at the planning station. Based on this model and imaging data, the surgical planning was completed. Following the planning phase, instrument setup was performed. The registration process was carried out by the operating surgeon. | NA |  | NA | NA | NA | NA | NA |  | NA | Na | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.1 | 3.66 | 1.44 | https://drive.google.com/open?id=1Ww04bqEi5KfmV7WIzSmmhBWE8mY5FDHT | https://drive.google.com/open?id=1nQXztaaM4A37x2tdlbh53S-palRMKlos | https://drive.google.com/open?id=1sGeOne5W1LXyOuUgKlO2S6_h4PNJfRIt
2026-03-17 10:27 | trushalgulhane1547@gmail.com | 2026-01-13 | Mahavir hospital | Surat | NA | Dr. Jenil | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | NA | Hospital: Mahavir Hospital, Surat Date: 13 January 2026 Procedure: Brain Biopsy Surgeons: Dr. Jenil and Dr. Hiran Modi Technology Used: Optical Navigation Technology Procedure Duration: Approximately 2 hours Pre-operative Details: Patient MRI data was provided by the operation theater staff prior to the procedure. The MRI data was successfully imported into the navigation system, and an accurate three-dimensional anatomical model was generated. This model was used effectively for surgical | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.2 | 3.53 | 0.78 | https://drive.google.com/open?id=15XQ2ptC_YI4tXYIHCdfrdfUJ-Tx1lnnj | https://drive.google.com/open?id=1PCjxS6nJfD-nZzj7c7vJk730aMWE7Fk6 | https://drive.google.com/open?id=1nhF3yVxfjdR71PPLA1f_7m1tPDikCgPb
2026-02-03 18:56 | gurjit@rxoom.com | 2026-02-02 | Aaihms Hospital | Surat | NA | Dr. Maulik Patel | Gurjit Singh | Cranial | Pituitary Tumor |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded, skull tracker with mini plate was applied. Patient registration was performed successfully, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and confirmation | NA |  | Strongly agree | Strongly Agree | Strongly agree | Liked the simple UI the most. | NA |  | One glyonce coating | NA | Gurjit Singh | Optical | 1.2.24 | Other | NS TRL 04 DB 101 | NA | 0.33 | 4.1 | 0.63 | https://drive.google.com/open?id=1KEYmyLFMRN1WgUNsJs-uJR5hHXerBCS4 | https://drive.google.com/open?id=1ILQ3EESPRBodvIAtpUJ_6f3fcfrD-otz, https://drive.google.com/open?id=150PmnST2ewOJulzNOMnLYdmxoGKNm_AX, https://drive.google.com/open?id=16EOsbJhzWDp2AfEtu3mfcqBJCrujyQad | https://drive.google.com/open?id=1Dq3RkSD6hZw6C8igJ8l0vOsUO7nrnsHs |  |  | https://drive.google.com/open?id=1oD3esQGCLJQaT3ezsXGzECAVDMyp4C0O
2026-03-12 20:33 | gurjit@rxoom.com | 2026-03-12 | Aaihms Hospital | Surat | No | Dr. Hitesh Chitrode | Gurjit Singh | CT-Spine | Spine Fixation |  | 1000 | Yes |  | Partially | Difficulty in registration and spine clamp movement while using hammer for pedicle screw. |  |  |  |  | NA | Pre operative CT scan was used and Dr. successful register single level L4 vertebrae after some tries of registration Surgeon confirmed the accuracy then accuracy lost after some hammering on vertebra. Then Dr. did multiple try on L5 but could not register vertebra properly and abort using the system . | Better spine reference and fast and easy registration method for spine. |  | Agree | Agree | Agree | More time consuming in spine and in brain surgery very useful. | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.24 | 3.46 | 1.06 | https://drive.google.com/open?id=1tGMR4xq0b9ajnnXXM82LyJ1wMpY5jrj8 | https://drive.google.com/open?id=1940f1Vz2EU_xq5eXkXGvIb86Svy75GDL | https://drive.google.com/open?id=1kDj0HIFEtaj9K271koSV3nwe-L84bJFE, https://drive.google.com/open?id=1hEssBzd5pGAW2MypXPzJEfHFT6Br-qDD | https://drive.google.com/open?id=1bPERIwFCrhmnLmzv1VVmMt3kyq5AI_mG |  | https://drive.google.com/open?id=1wdJkF-RonIa15jm-1hY08gKsUj6Eodwo
2026-03-14 17:53 | gurjit@rxoom.com | 2026-03-14 | Aaihms Hospital | Surat | NA | Dr. Maulik Patel | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brian Biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and Skull tracker with mini plate was applied and registration was done successfully by surgeon, and accuracy was verified and accepted by the surgeon. Bur hole was made and surgeon aligned the biopsy guidance tool with pre planned trajectory and after applying the stopper on measured distance, surgeon took biopsies. Navigation was used completely throughout the surgery for guidance. | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | Liked the Simple Registration process and UI. | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.2 | 2.03 | 1.01 | https://drive.google.com/open?id=1U4KkZFWidNkRF2mtjin2Fl3jPsXKJ2SZ | https://drive.google.com/open?id=1EBbZ7mD5_6ruz8nA02V7SI79lxK9uGx9 | https://drive.google.com/open?id=1nkxR4q-9UHsl-ab6vr3cJUw-E1Hq9rFD, https://drive.google.com/open?id=1xrHVja23DGhZ99y15grFJ2fGVOxZGrGl, https://drive.google.com/open?id=1vMryruzyr3LlBt2SbopUJ5FfOuT4zihT |  |  | https://drive.google.com/open?id=1IlA-ZbYhw5XY46dUQqwqg5z1oebh4558
2026-03-17 10:33 | trushalgulhane1547@gmail.com | 2026-03-11 | Sindhu hospital | Hyderabad | NA | Dr. Narsimbharao | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | *Surgery Report* Technology: Optical Technology Date: 11 March 2026 Hospital: Sindhu Hospital, Hyderabad Procedure: Awake Tumor Excision Surgeon: Dr. Narasimharao Navigation Accuracy: 0.79 mm Outcome: Successful The procedure began with the utilization of MRI and T2 imaging data provided by the OT staff. The imaging data was imported into the navigation system and processed at the planning workstation to generate a detailed 3D anatomical model. Based on this model and imaging dataset, the surgical planning was carefully performed to guide the procedure. Following the planning stage, the instrument setup was completed and the patient registration process was performed by me to align the patient anatomy with the navigation system. The system maintained a stable accuracy of 0.79 mm throughout the entire procedure without any disruption, ensuring precise guidance during the surgery. The surgery was completed successfully. Post-procedure, a discussion with Dr. Narasimharao confirmed his satisfaction with the neuro-navigation system, particularly appreciating its accuracy and stability during the operation. He also expressed interest in utilizing the system for future procedures such as DBS (Deep Brain Stimulation) and frame-based biopsy. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.22 | 2.48 | 0.79 | https://drive.google.com/open?id=1ENjHLnMgybTzhnomMNKEv7cpJAlzUKOI | https://drive.google.com/open?id=10IrB6WS094brFtiFBCNnghuIBdINbPBQ | https://drive.google.com/open?id=1eTawQLt37xBeaAOvWuMwDnxCJ-0h4P_9
2026-03-17 13:50 | gurjit@rxoom.com | 2026-03-17 | Aaihms Hospital | Surat | NA | Dr. Maulik Patel | Gurjit Singh | Cranial | TNTS Pituitary |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and Skull tracker with mini plate was applied and registration was done successfully by surgeon, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance. | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | NA | NA |  | One Glyonce | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.21 | 4.32 | 0.57 | https://drive.google.com/open?id=1kRmgltB5qbI0-guqwmlTpMHgwHmbkLCD | https://drive.google.com/open?id=1vBOVcfXoFy5TKNSCp8fkzi9_BNRp7CEX, https://drive.google.com/open?id=1oXdvQ6pB6MQRwjm7NcngrDFYz6JUw52G, https://drive.google.com/open?id=1SnDSiB92L6fPP70QqkUGcpB-hOEPMpWZ, https://drive.google.com/open?id=1qcmSrABaW6aXau_tN00PPzkKTN3xO7l8 | https://drive.google.com/open?id=1WaktR5pAOG-XWrnMiHXNpbeMf_NmvHNg |  |  | https://drive.google.com/open?id=1kY8ZfXAHhyYLFCZxaQLsyT35MqgxBgWZ
2026-03-18 16:33 | gurjit@rxoom.com | 2026-03-18 | Unique Hospital | Surat |  | Dr. Harin Modi | Gurjit Singh | CT-Spine | L4-L5 Fixation |  | 1000 | Yes |  | Not used | Registering was done on single level vertebra and Surgeon checked the accuracy by placing the probe on pedicle entry point then dr reported that it is not accurate and it it is showing on S1 instead of L5 Pedicle entry point. Dr did the registration again but same issue reported by surgeon and system was not used. |  |  |  |  | NA | Pre operation CT scan was loaded and surgeon did the single level vertebrae registration on L5. Registration was done complete surgeon confirming the accuracy but Dr reported it is showing incorrect. I suggested the reregistration and did screen recording for finding out what is the issue. Again registration was completed but surgeon again reported that I am keeping probe on L5 Pedicle but it is showing on S1. Then put the k wire in L5 and took the X-ray shot then put the probe in it leave it. In navigation it was showing on S1. I asked surgeon to check the accuracy on points where we did the registration surgeon reported on those place it is showing correct but then again on putting on L5 Pedicle it showed in S1 then surgeon stopped using the system. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.11 | 4.63 | 0.95 | https://drive.google.com/open?id=1ROsrbN-yzeNok2B_pwnTolnDKHeiP36r | https://drive.google.com/open?id=1rIlDLWQl0LMaoNxaVPUCXhDH8767CF45, https://drive.google.com/open?id=1Wcx3ay_ty5Kb8ehsS0BS-z_GtusMt0mW | https://drive.google.com/open?id=1-CjTEoH_KbaZrA1iScuHRrRz1aIUmDXx, https://drive.google.com/open?id=1aMOYln01sZMQFzJwi2Eq-kC80ZUZrX6q, https://drive.google.com/open?id=1-qSPSwcyA-H6uO4-rH7G_lgt3aRVduDL, https://drive.google.com/open?id=12GAKZxLc_gYTwXxh-noFU6V-hpgZAfFD
2026-03-18 16:46 | gurjit@rxoom.com | 2026-03-18 | Aaihms Hospital | Surat |  | Dr. Maulik Patel | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brain Biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and Skull tracker with mini plate was applied and registration was done successfully by surgeon, and accuracy was verified and accepted by the surgeon. Two Bur hole was made and surgeon aligned the biopsy guidance tool with pre planned trajectory and after applying the stopper on measured distance, surgeon took biopsies from two different targets. Navigation was used completely throughout the surgery for guidance. | NA |  | Agree | Agree | Agree | Agree | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 100 | 4.72 | 0.57 | https://drive.google.com/open?id=1_XvriMv_Un40sWUc9PwNbsRCHOfZnxRo | https://drive.google.com/open?id=1cetSwXhhVaNi5DsVyZ37uPjN1TQZOCLw, https://drive.google.com/open?id=1VUGk46mojjfrIOsuiqeiuTzi-LmSydqE | https://drive.google.com/open?id=1AU-WIQZ_BLfz_hxuK0eqQYuaEnLlEsEM, https://drive.google.com/open?id=1hm4jbUHjbnKeLsyk1HHJwkd-XOVF_Kq_
2026-03-19 19:12 | gurjit@rxoom.com | 2026-03-19 | Aaihms Hospital | Surat |  | Dr. Maulik Patel | Gurjit Singh | Cranial | Brain Tumour |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and Skull tracker with mini plate was applied and registration was done successfully by surgeon, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance. | NA |  | Agree | Agree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.26 | 4.63 | 0.56 | https://drive.google.com/open?id=19pzFuajvxluQrgxcPmmwONyVrPjxfmGY | https://drive.google.com/open?id=1q8z979eGSS-RHDgsucoT6Z3fm65qgLCg, https://drive.google.com/open?id=14Ub9LZWGQRBbRZKGz53neGxsl3M_P6rq, https://drive.google.com/open?id=1lnzQSp7XjcfQY0ha0Mcs5KLDS4ekKN8W | https://drive.google.com/open?id=1jFQyGlJ_cYVdzTIrR50nOppcHuzFn924, https://drive.google.com/open?id=1d8_cB9xmba9icCUe6-PIkKi7IsdIrarQ
2026-03-23 22:41 | gurjit@rxoom.com | 2026-03-23 | Avdhoot Hospital | Vadodara |  | DR.Yogendra Bhatt | Gurjit Singh | Cranial | Aneurysm Clipping |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CTA scan was loaded and Standard reference with clamp adapter arm was applied and registration was done successfully by surgeon, and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance. | NA |  | Agree | Agree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.17 | 4.49 | 0.3 | https://drive.google.com/open?id=1ogB41bMn6f5G6lavNAghYZo5xdDr0oNI | https://drive.google.com/open?id=1x4tXXgbL8V_D6CG4k6WV3CH-q0d03vvn, https://drive.google.com/open?id=1txPSUS8OkW6ES1rD4K714Cp1twOTtaYP | https://drive.google.com/open?id=11_9evuh-x27zmE83LcPdRN0Ots-FqS9R, https://drive.google.com/open?id=15xZvymWLySR0eJNxds3qCapnhnO4_G30, https://drive.google.com/open?id=1oridPuJbNi4rvsiB06wi78eG4LbXQJNP | https://drive.google.com/open?id=1XH0hl5PabRoyKvG3cUZKKUMmlhro_kBK | https://drive.google.com/open?id=1wcYdOTkHs8bF-hIEOc84f5k2PcAc5fRo | https://drive.google.com/open?id=1RnfQGMMpnYtpBWiLojq0VQSgXQ5QBlJv
2026-03-25 21:46 | gurjit@rxoom.com | 2026-03-25 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | C1- C2 Fixation |  | 10 | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded and CT clamp reference was applied and single level vertebra registration was done successfully by surgeon for both C1and C2 right pedicle , and accuracy was verified and accepted by the surgeon. Navigation was used completely throughout the surgery for guidance and both screw was placed with the guidance of navigation after placing screw confirmation was done with C-arm. Due to some complications in left side vertebrae screw was not place. | NA |  | Agree | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.29 | 2.68 | 0.09 | https://drive.google.com/open?id=1pdrQvgh1y42V9DkGV5b7nNjAS4vgz88v | https://drive.google.com/open?id=1sn8i3XrnjJgMcv3-3wydwG9zyiqfKViw, https://drive.google.com/open?id=1kZuPHfgVAonUGGm1cQjBj6wTJyIAxLw_, https://drive.google.com/open?id=17bg3cijrmMgTT1yjjmNLIRjSNQoYMWyF | https://drive.google.com/open?id=1XEGuxkMhfercTF76cdEd5wutapiuRfTs, https://drive.google.com/open?id=10rWmcLVDX_yMxU26cWX0HiRThrvZ9r08, https://drive.google.com/open?id=1KbGcMo9MBnhHpfD0HWtotDLpNvnTTOrJ, https://drive.google.com/open?id=1z5kMmjToUqmBeG27sevPyUMTAqfKphAE, https://drive.google.com/open?id=1iPob2DRleZSAKSHO6sNHatf10OjPab-Y
2026-03-28 00:36 | gurjit@rxoom.com | 2026-03-27 | Avdhoot Hospital | Vadodara |  | Dr.Yogendra Bhatt | Gurjit Singh | CT-Spine | C1-C2 |  | 1000 | Yes |  | Partially | [5 empty] | NA | Preoperative CT scan was loaded and CT clamp reference was applied and single level vertebra registration was done successfully by surgeon for C2 , and accuracy was verified and accepted by the surgeon. Navigation was used one screw was place with guidance and confirmed with C arm. After Clamp was moved during the surgery and surgeon did the registration and surgeon verified the accuracy. In the end surgeon express less value addition of navigation in spine. | NA |  | Neutral | Better Spine clamp | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.12 | 2.82 | 0.27 | https://drive.google.com/open?id=1GFnn7Gttnua9V7FcYVWJU7mTGhNKk0Qg | https://drive.google.com/open?id=1H7DzriI3QYylmE-lf7dqPdIJkyS7Ch-e | https://drive.google.com/open?id=1-DS2yG_EvHTF8wI5in1V_qcNUbo8xRg7, https://drive.google.com/open?id=1R8uoAdVJCcFGs3GcNZ233sq3YheP89Wg, https://drive.google.com/open?id=11hjcwHHQSumydxnimjCfmveH2bY5HNfc, https://drive.google.com/open?id=16u12X7vhqdAWJIsMbB-8MGxQind3NGmc
2026-04-08 15:49 | gurjit@rxoom.com | 2026-04-08 | KIMS(PMSSY) | Hubli |  | Dr. Sateesh R | Gurjit Singh | Cranial | Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT and MRI scan was loaded and Standard reference with clamp adapter arm was applied and registration was done successfully, and accuracy was verified and accepted by the surgeon. Tumor was at surface level of brain. Surgeon used Navigation for craniotomy marking. After craniotomy surgeon didn’t found tumor at the center but beside of it. Navigation was correct on tumor. Dr. understand but asked me to corrected him for craniotomy. | NA |  | NA | NA | NA | NA | NA |  | One glyonce | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.11 | 3.24 | 0.51 | https://drive.google.com/open?id=1I9Lp1QfMTSIbTRQNutmwe9nqqxp_4ZYX | https://drive.google.com/open?id=1VW5t2v7lx_QkhC7exqpm9qCMEez_6c1x, https://drive.google.com/open?id=1BxqAfmLcDBXvtMdqumVIq_cQfChDzq67, https://drive.google.com/open?id=1SdH-QORVXWe42zeNOOlmWtTRBnBhZ97R | https://drive.google.com/open?id=1z81KcTLsBc9zh91KjEMGot7TV56nkfXb, https://drive.google.com/open?id=1a-LpeEBsjaTcHV-9nZpxS9i9cy4MrzbH
2026-04-13 13:10 | gurjit@rxoom.com | 2026-04-13 | Kims(PMSSY) Hospital | Hubli |  | Dr. Vishawanatha N | Gurjit Singh | CT-Spine | Spine Fixation D12-L3 |  | 1000 | Yes |  | Partially | [5 empty] | NA | Preoperative CT scan was loaded and CT clamp reference was applied and single level vertebra registration was done successfully by surgeon for L3 , and accuracy was verified and accepted by the surgeon. Navigation was used for both side pedicle finding and navigable tap and pedicle finder was used and confirmed with C arm. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.2 | 3.45 | 0.24 | https://drive.google.com/open?id=1Ex-mzIg8h6Gpr3lapquefxINISOGnM1p | https://drive.google.com/open?id=1NICNG_MN8ubiXltb_Wy-7joY-2xLXgVr | https://drive.google.com/open?id=1yrbHNQRNXKSIXEsjQEt5f4SdXkXs-Wup, https://drive.google.com/open?id=1Vifjk8tDpo2oSsPlQrmkNmczIx3QuHPt, https://drive.google.com/open?id=1nOD8yFZazC4qHh6JNxJir-ocRaPK_Iju, https://drive.google.com/open?id=19rp6BQoD37B9CbGJ3W40ycEEiX9IU5ID, https://drive.google.com/open?id=135LT8DtaSpD9OKyZQc6dGaFJYY7z8P5W
2026-04-15 19:14 | gurjit@rxoom.com | 2026-04-15 | KIMS(PMSSY) | Hubli |  | Dr. Sateesh | Gurjit Singh | CT-Spine | C1-C2 Fixation |  | 1000 | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded and CT clamp reference was applied and single level vertebra registration was done successfully by surgeon for C1- C2 , and accuracy was verified and accepted by the surgeon. Navigation was used for both side pedicle finding and screw trajectory guidance. Surgery was getting oriented to navigation guidance than previous case. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.2.24 | NS TRL 02 MF 201 |  | NA | 0.16 | 1.74 | 0.32 | https://drive.google.com/open?id=16UHfqDl9NgJSixKUZsjz3gU87pTmylwc | https://drive.google.com/open?id=1YAggplJ6DZ-7sRxi76jDgyy8sPXqT3rD, https://drive.google.com/open?id=1-crpIgcTQdhxvjjw_HM6VePkqRDwn1BZ, https://drive.google.com/open?id=18pQ5rOVgxbat7Uh154EGhhVI-2IfCCiz | https://drive.google.com/open?id=1Qd_YY6CqG-ZJPg6hnguUanT3Qq1QH23F, https://drive.google.com/open?id=1XLTfZa3_Un3g-UevS5gggqAWJSLo-OS4, https://drive.google.com/open?id=1kLIRbKhv35wojwiwHF7_SAd5e9DYRQWA, https://drive.google.com/open?id=1DMe4qlsKf83v7u6xyA_6YDZKZhyzQvXe | https://drive.google.com/open?id=1Z2nUZPF-DZyAbVlwf7xGtPWKD9g5vjrd |  | https://drive.google.com/open?id=1VLZc1-hgSFgrC6VSp4Y0IKzymA7EyUD9
2026-05-16 18:18 | trushalgulhane1547@gmail.com | 2026-05-16 | Bharati hospital | Pune |  | Dr. Anand Doshi | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other |  | NA | 0.13 | 4.29 | 0.88 | https://drive.google.com/open?id=1uh8fQX-bRXWMpXQtioj9PpZmX_72pNlA | https://drive.google.com/open?id=1PNcMKNQGgVhyYe3qzXnYzlLABRM8f4gy | https://drive.google.com/open?id=1DwH0S7LNk5yz8h14yuZEe7pLWZFzkyX_
2026-05-02 15:30 | gurjit@rxoom.com | 2026-05-02 | BMHRC | Bhopal |  | Dr. S Dixit, Dr. S Gupta, Dr. Sandeep | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brain Biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI was loaded. A brief system introduction was give to Surgeon and his team from how to start, load and what are the main functions software. Junior surgeon marked the trajectory and patient was registered by surgeon sand accuracy confirmed and accepted by surgeon. A burhole was made and biopsy trajectory was fixed using biopsy guidance tool and biopsy was taken using navigation guidance. | Provide a needle tracker tool according to their specification. Create a Quick steps reference document for how to start system and load and what’s main function of software and how to use them. A sleeve for biopsy guidance nozzle to rectify play and snug fitting of biopsy needle. |  | NA | NA | NA | NA | Relevant length marking on tools like biopsy guidance tool nozzle and sleeve, spine tools. |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | Other | NS TRL 02 MF 208 | NA | 0.06 | 4.82 | 0.87 | https://drive.google.com/open?id=1Ekykpzm56rqnzTI68NqiyQhHcvWX01YF | https://drive.google.com/open?id=1a-Guf-y4otmeKzJF-bM-oJ7K3deLW5Op, https://drive.google.com/open?id=1hELnPXK--XeYad13I-m2zsKlV3TTCZnT, https://drive.google.com/open?id=1b6dqZQFQdu1tNw02Q31-eL2air-Ukce2 | https://drive.google.com/open?id=1qL566A6AZk_hAqCFF8eUWhNzoL5Gta1r, https://drive.google.com/open?id=19SFB4qz1apo-C0NABa70AGIBTw3PJ3nz
2026-05-04 16:35 | gurjit@rxoom.com | 2026-05-04 | BMHRC | Bhopal |  | Dr. S Gupta, Dr. S Dixit | Gurjit Singh | Cranial | Tumour excision |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT was loaded and standard reference was applied on skull clamp adapter. Patient was positioned in prone. Fiducial marker Vitamin E method was used on patient. Registration was done by and accuracy accepted by surgeon. Navigation guidance was used to craniotomy and trajectory planning. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | Other | NS TRL 02 MF 208 | NA | 0.11 | 3.8 | 0.58 | https://drive.google.com/open?id=1cJOgYcQDoN0TwyPVY7EOF18SIH1xt-j6 | https://drive.google.com/open?id=1AzD9MifwBWr9jNmB2C7sEPRV2W54HBNx, https://drive.google.com/open?id=15Tz6rGCwQ10pS0qd57WKHLRf9m5mSy8w | https://drive.google.com/open?id=16xuofcwiFM3-Z9_iCmqy3ZajJQjZKF-1, https://drive.google.com/open?id=1u3AUak565-bixFND9gasHpH9zmnURGx8, https://drive.google.com/open?id=1c54tJevVrSMCMpuX-eoPgQE8GEdRTi2d
2026-05-06 14:20 | gurjit@rxoom.com | 2026-05-06 | BMHRC | Bhopal |  | Dr. S Dixit, Dr. S Gupta | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brain Biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT was loaded. Junior surgeon marked the trajectory and patient was registered, accuracy confirmed and accepted by surgeon. A bur-hole was made and biopsy trajectory was fixed using biopsy guidance tool and biopsy was taken using navigation guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | Other | NS TRL 02 MF 208 | NA | 0.31 | 6.21 | 0.52 | https://drive.google.com/open?id=1sojTpw1636LxsRYa3mjK805lx_SEln2s | https://drive.google.com/open?id=1Qa4coF-hZwtwhLSCSpwhCe-7a4ALH2tu, https://drive.google.com/open?id=1Enm3lzLeCAi7ESESkbZUmJNCjI6XgoDW | https://drive.google.com/open?id=1NVsqYRItnFP1-j2FmAUXEQ1_tOnx67rr, https://drive.google.com/open?id=1ef9cxwHIFwa-sHK7XtklhiW4V1FcDcUS
2026-05-09 15:43 | gurjit@rxoom.com | 2026-05-09 | BMHRC | Bhopal |  | Dr. Sandeep Sorte, Dr. S Dixit, Dr. S Gupta | Gurjit Singh | Cranial | TNTS Pituitary tumour |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded. Skull tracker was applied using our demo screw driver and screw set. Registration was done, accuracy was tested and accepted by Dr. Sandeep Sorte. Navigation system was used during surgery for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | Other | NS TRL 02 MF 208 | NA | 0.14 | 2.71 | 0.51 | https://drive.google.com/open?id=18OZUyI02s7XuX36ymUzoJh4l_bgPsqHY | https://drive.google.com/open?id=1cBm1sxWjgH8DuWMmkxjelYrRqZRi6zun | https://drive.google.com/open?id=1rP3AQOSgr0nal-VKJAnbKQSUK07gdZee, https://drive.google.com/open?id=1XF98EQkPSLnI5Dmf7rDH7bgWCACB2CxZ
2026-05-09 19:20 | trushalgulhane1547@gmail.com | 2026-05-09 | DMH Hospital | Pune | NA | Dr. Tushar borde | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | Biopsy tracker with sleeve need to be provided | Successful biopsy case | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.26 | Other | DMH hospital | NA | 0 | 2.68 | 1.11 | https://drive.google.com/open?id=1mpbwp-qf0LsWLasBgRG_X5sq0uMQ21ED | https://drive.google.com/open?id=1zxtHopesX_5yq5mDlY2RsxifIqvdSrWJ | https://drive.google.com/open?id=1n4MVgKJn9S8kXwuLuN_7hMvAl1xa1jC3
2026-05-24 19:33 | trushalgulhane1547@gmail.com | 2026-05-23 | Nizam Hospital | Hyderabad | NA | Dr. Vamsi Krishna | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | Trajectory planning editing in probe eye view. Entry point button in track plan | Trajectory planning editing in probe eye view. Entry point button in track plan | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system | Nizam Hospital Hyderabad | NA | 0.7 | 2.9 | 1.01 | https://drive.google.com/open?id=1_tvC1MRxUeFRKlo7jiUNCnnvcGfrcp_W | https://drive.google.com/open?id=1gpgX-kMReJuQqYHkgVD2NPwtAySHmUx_ | https://drive.google.com/open?id=1nrozdjv-47f1CfA6Q6cqI1-H9Lk-NPw8 |  |  |  |  | YES, this demo is at an existing customer location.
2026-05-30 21:12 | trushalgulhane1547@gmail.com | 2026-05-29 | Nizam Hospital | Hyderabad | NA | Dr. Nikhil, Dr. Murli, Dr. Prajwal | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful case | NA |  | NA | NA | NA | NA | NA |  | One 10mm screw head got damage | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | 2 eight mm and 2 ten mm screen and screw driver | 0.16 | 2.29 | 0.66 | https://drive.google.com/open?id=1E-Mta-Q53yaaWMNtUR00-vxOXKnUwU8F | https://drive.google.com/open?id=1mO9JmAa0gQlYhyJXmiYL3g6UOkdyg4R8 | https://drive.google.com/open?id=15RoAabopLZoBJOz6Thno-C0fd-4Hys1X |  |  |  |  | YES, this demo is at an existing customer location.
2026-05-30 21:17 | trushalgulhane1547@gmail.com | 2026-05-30 | Nizam Hospital | Hyderabad |  | Dr. Vamsi Krishna | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.2 | 4.46 | 1.5 | https://drive.google.com/open?id=13I5G1sUcOGZiJTTkdvzu4pYLIimw62jS | https://drive.google.com/open?id=1ZvfUuUtRa72P0wNiXQ_JTCGyoty45BRa | https://drive.google.com/open?id=1FodaEcdFHhTXW0LixT90BlFb50gj93We |  |  |  |  | YES, this demo is at an existing customer location.
2026-05-30 21:21 | trushalgulhane1547@gmail.com | 2026-05-25 | Nizam hospital | Hyderabad |  | Dr. Vamsi Krishna | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.7 | 4.67 | 0.58 | https://drive.google.com/open?id=1MUoVX1xXgdtzijG16To1MSRruMBVbrDV | https://drive.google.com/open?id=1n9HsRRVEq93XEhIlcOlYMh9PCmPhqxjs | https://drive.google.com/open?id=1_EbGPLHdgx2RXoXOcggLqROytnx4VFkM |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-02 21:20 | hemarshitha1802@gmail.com | 2026-06-02 | Jalappa hospital | Kolar | NA | Dr Venkataramana | Hemarshitha H S | Cranial | Left RMSO craniotomy & excision of tumor |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.28 | Other: Please write Serial Number of the system | NX01OP003 | NA | 0.28 | 3.94 | 1.14 | https://drive.google.com/open?id=1EFVgI006WA81EqhrmnzhkmL-8xCKU1dS | https://drive.google.com/open?id=1CusVQSKAd9InuWmUjlm26l74NwWhN6I9 | https://drive.google.com/open?id=1qAsF_Yjv99_cfoVhdkvLminqiXXyoj90 |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-04 12:20 | trushalgulhane1547@gmail.com | 2026-06-02 | Nizam hospital | Hyderabad | NA | Dr.murli, dr.prajw | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.2 | 2.74 | 0.67 | https://drive.google.com/open?id=1i2csmiGBYQcN_E_B1OKb01rabyby4tmD | https://drive.google.com/open?id=1sPE8-MbGJ6Bi95pMjlxiBlBIqglJZmCv | https://drive.google.com/open?id=1UXuEI98USMxo_3UAmF61sWXCJO9Nd541 |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-04 12:39 | trushalgulhane1547@gmail.com | 2026-06-02 | Nizam hospital | Hyderabad |  | Dr. Thota satya swaroop | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.7 | 4 | 1 | https://drive.google.com/open?id=15kAOZ3Ltp0M89lx-yiWfU-j34YdnWqbS | https://drive.google.com/open?id=1dma7x1g48jdNCVjX4TAmX2j0Qv4SoEjW | https://drive.google.com/open?id=1tVfeZcwdQy9_arE88hOXNG9fv_J1pjAg |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-04 14:21 | trushalgulhane1547@gmail.com | 2026-06-03 | Nizam hospital | Hyderabad |  | Dr. Suchinda mam | Trushal Gulhane | Cranial |  | C1,C2 through transnasal |  | No | Depth deviation after drill near the C1 bone | Not used | [5 empty] | NA | Surgery Report Date: 3 June 2026 Hospital: Nizam Hospital, Hyderabad Surgeon: Dr. Suchinda and Team Case: transnasal c1, c2 Navigation Accuracy: 1.58 mm Pre-Operative Details: The patient imaging data was provided by the OT staff and loaded into the navigation system. Procedure Details: Surgical planning, skull reference fixation, instrumentation setup, and patient registration were completed successfully. During accuracy verification, the navigation accuracy was found to be good and acceptable by the surgeon. The surgery was performed through a transnasal approach to access the C1 and C2 region. During the procedure, drilling was carried out on the anterior part of the C1 and C2 bones. After drilling, when the probe was placed on the bone surface, the navigation system showed the probe slightly above the actual bone surface. This was most likely due to anatomical or bony shift caused by drilling during the procedure. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.7 | 3.35 | 1.56 | https://drive.google.com/open?id=1DJ3HX_gZW5J72G1QEHrZoF3FmH2VrJku | https://drive.google.com/open?id=1F5DGgNU9RFg6NbqgXJq6DZDXbSOawfXI | https://drive.google.com/open?id=13w-nasfb4mhk_xCCCiWHUVXLEHvk1lcG |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-04 22:38 | trushalgulhane1547@gmail.com | 2026-06-04 | Nizam hospital Hyderabad | Hyderabad |  | Dr. Vamsi Krishna | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration | NA |  | Good | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | EM | V1.3.28 | Other: Please write Serial Number of the system | EM system | NA | 0.26 | 2 | 1.15 | https://drive.google.com/open?id=1MajkOzp1cHkrveqapfG6RyZCPjiVrY2x | https://drive.google.com/open?id=1ZsGexF9lKPS_lCj0PznLsFW81rPB4hcg | https://drive.google.com/open?id=1EF_NXLDrAplrNbmQ7EOl93LX0Fc4VY13 |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-09 12:01 | trushalgulhane1547@gmail.com | 2026-06-08 | ESI hospital | Hyderabad |  | Dr. P. Rajesh | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of neuro navigation system | NA |  | Good accuracy | Good | Good | NA | NA |  | NA | Surgeon said for previous case leyla arm was not stable | Hyderabad Distributor | Optical | V.2.24 | Other: Please write Serial Number of the system | Hyderabad demo system | NA | 0.2 | 4.75 | 0.5 | https://drive.google.com/open?id=1serZ6tHvfDB0PJ7GZTGrzM_WA1zm1vEP | https://drive.google.com/open?id=13GStEPmcH8YyesxhtbpX7MTix3lN-NSD | https://drive.google.com/open?id=1w-ePJXu8sSSpykn9bBGcmHEJskSyo_E4 | https://drive.google.com/open?id=1DNRijYEFbX4qwrLhutD1UOVCfSVkWGbS |  | https://drive.google.com/open?id=1DYn69S0gIaGtZzCUa4IDpt1F7aFQya_Y |  | NO, this demo is not at an existing customer location.
2026-06-09 12:09 | trushalgulhane1547@gmail.com | 2026-06-08 | Nizam hospital | Hyderabad |  | Dr. Thota satya swaroop | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful case | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.02 | 3.56 | 0.58 | https://drive.google.com/open?id=1CM9BjDzUNRsRzTaxF-of1d2zCOgky0af | https://drive.google.com/open?id=1wwvQY8uT4eUXX7boT8sQnQMiKi2ufun2 | https://drive.google.com/open?id=1lVKiJLgKDxTZ9CCgWLQ17PaOmiLGdLwe |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-09 20:00 | gurjit@rxoom.com | 2026-06-09 | Hinduja Hospital | Mumbai |  | Dr. Ujwal Yeole | Gurjit Singh | Cranial | Endo lymphatic sac tumor |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT and MRI scan was loaded and merged. Standard reference was used and registration was done, accuracy was confirmed and accepted by surgeon. Navigation was used during the surgery for its guidance. | NA |  | NA | NS | NA | Fast registration | NA |  | NA | NA | Mumbai Distributor | Optical | 1.2.24 | Other: Please write Serial Number of the system | Hinduja Hospital, Mumbai | NA | 0.12 | 2.73 | 0.7 | https://drive.google.com/open?id=15w2j3_hK08HF1WvI2FoBnG0Xp8dbj_ss | https://drive.google.com/open?id=11NdJHKHU3b8NPe7gLJde1sKKixkv5gMu | https://drive.google.com/open?id=1s3_YlUySXlIgto_Mpq6YkmXwzXiCs2KO, https://drive.google.com/open?id=138lTCfJ7uP1bVRO-7Eg-un6LjUHF8rSC |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-09 22:19 | trushalgulhane1547@gmail.com | 2026-06-09 | ESI hospital | Hyderabad |  | Dr. P. Rajesh and Dr. Chandrakant | Trushal Gulhane | CT-Spine |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of spine navigation | NA | NA | Good | Good | Good | NA | NA |  | NA | NA | Hyderabad Distributor | Optical | V1.2.24 | Other: Please write Serial Number of the system | Hyderabad demo system | NA | 0.16 | 2.53 | 0.82 | https://drive.google.com/open?id=1iAu3quJG0M7SBmL8oFyWledpy1pw2-8j | https://drive.google.com/open?id=1VJtGyhq6FdcnY3BhcLDg3dzFcr9aosck | https://drive.google.com/open?id=13i9nPLmsyCtofox9EJzAqZaRVQWu0oNj | https://drive.google.com/open?id=1Gd4gQ16xjByQrZr9naPIDWfhQnk4FGXi |  | https://drive.google.com/open?id=1zwBtceOoOdGq1Y_wbNxH51mXppNBgSxM |  | NO, this demo is not at an existing customer location.
2026-06-10 20:52 | trushalgulhane1547@gmail.com | 2026-06-10 | ESI hospital | Hyderabad |  | Dr. Chandrakant | Trushal Gulhane | CT-Spine |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of spine navigation system | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Hyderabad Distributor | Optical | V1.2.24 | Other: Please write Serial Number of the system | Hyderabad demo system | NA | 0.16 | 1.69 | 0.29 | https://drive.google.com/open?id=17foOR85XiVSyxSwXHiqgNGOYS3K5LDwV | https://drive.google.com/open?id=1--ncpsN5KXB4OJA1yaGFBZ3GO1Gsmv7_ | https://drive.google.com/open?id=1UfdY6SwQmFgzgIxTYJyFI7oQRjz1JvM7 | https://drive.google.com/open?id=1Pr883TbUrnhbpRklDQKuyHitZEtoHGJN |  | https://drive.google.com/open?id=1ii09V1o5nUA_pX92a-1teZXtfHjF4y1W |  | NO, this demo is not at an existing customer location.
2026-06-10 20:14 | hemarshitha1802@gmail.com | 2026-06-10 | Vydehi hospital | Bangalore | NA | Dr Ananth Kishan & Dr Benak | Hemarshitha H S | Cranial | Biopsy |  |  | Yes |  | Partially |  |  |  |  | The surgeons marked the planned craniotomy points using the navigation system. After the bone flap was opened, a discrepancy of approximately 5 mm was observed between the actual anatomy and the previously marked craniotomy points | NA | Need to analyse the problem of accuracy occured in this demo | NA |  | There was almost 5mm difference in the tumour location | NA | Ease of use | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.28 | NS OCR V01 A4 005 |  | NA | 0.4 | 4.39 | 0.5 | https://drive.google.com/open?id=15rGL9t6_23mcYjjg-s6bhVWHXhKUiHfY | https://drive.google.com/open?id=1UY2ZMiaT4BbyEqSTmJUPgCqkQH877yVZ | https://drive.google.com/open?id=1iRQ0tRe3WD36CVWvNJW9OjkVYD6e8Ls_ |  |  | https://drive.google.com/open?id=1hYilzhEwWYkL4nAkLT-XgySY9x4NMunS |  | NO, this demo is not at an existing customer location.
2026-06-11 19:39 | trushalgulhane1547@gmail.com | 2026-06-11 | Nizam hospital | Hyderabad |  | Dr. Anil ,dr. Sanjay ,dr.swaroop | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of electromagnetic navigation | NA |  | Good | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.2.24 | Other: Please write Serial Number of the system | Electromagnetic Navigation | NA | 0.19 | 0.72 | 0.54 | https://drive.google.com/open?id=17eK5bDuXWxIH2K3cohsGJXYS2H61J3aX | https://drive.google.com/open?id=1geXIZDsW_hQMeDV8PI5C94zVp2vW6d8L | https://drive.google.com/open?id=1yx_kS2ZFQGLxZpG7aHP0GNn55vmUsIWd |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-13 21:26 | trushalgulhane1547@gmail.com | 2026-06-13 | Nizam hospital | Hyderabad |  | Dr. Ramanath and dr bhavana | Trushal Gulhane | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Partially | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.38 | Other: Please write Serial Number of the system | Nizam hospital | NA | 0.05 | 5.92 | 1.4 | https://drive.google.com/open?id=1X94-S-4ElPALy8fxKlIBeLetbdE8icbE | https://drive.google.com/open?id=1JcvfMFbh5FY-URSsGrYir0bhQDbihf4s | https://drive.google.com/open?id=16SVetnd3mSaff32N8wZBS88IgdQiYCKZ |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-13 21:58 | trushalgulhane1547@gmail.com | 2026-06-13 | Nizam hospital | Hyderabad |  | Dr. Akshay rajput and dr. Vamsi Krishna | Trushal Gulhane | Other | Stereotactic DBS |  |  | Yes |  | Not used | [5 empty] | Providing to points to see merging of CT and MRI | Dry demo of stereotactic DBS | Providing to points to see merging of CT and MRI |  | NA | NA | NA | NA | NA |  | NA | NA | Trushal Gulhane | Optical | V1.3.28 | Other: Please write Serial Number of the system | MacBook | NA | 1 | 1 | 1 | https://drive.google.com/open?id=1eavb1WLb6qln_bMtREVLTrlbAE5Tmdmm | https://drive.google.com/open?id=18D-KE5PfK09cSsJ5wYC-iTKxv2YKk3nz | https://drive.google.com/open?id=1Zie2ns4SGsB2V4TKtlN9X23jdGJTCskh |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-15 18:24 | gurjit@rxoom.com | 2026-06-15 | Aaihms Hospital | Surat |  | Dr. Maulik Patel | Gurjit | ENT | TNTS- Pituitary Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI and CT scan was loaded and merged. Skull tracker was used as a reference. Registration was done and accuracy was confirmed by surgeon. Navigation was used for its guidance during the surgery. | Surgeon wants an option to erase/correct the ROI incase marking go wrong, instead of discarding the whole marking/highlighting and start again for a big region. |  | Agree | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.14 | 6.8 | 0.51 | https://drive.google.com/open?id=1zjKVatpvyL_funyLhgJTv4YHeTLXmCqV | https://drive.google.com/open?id=1gm-SbOdBSXWED8XLwsuQ05Ld_eK4RciC, https://drive.google.com/open?id=1Qu0_BOReoc2VoTCMTzeWSxMHziGGamyv | https://drive.google.com/open?id=1cJY-PJjZWEo8J6FHbd_tQatKpicUdaLf, https://drive.google.com/open?id=1QdagQBBW8-8pTx__4ECMgbjhWijSXV0Z, https://drive.google.com/open?id=1ELDGaYffHYKg3dXm0plpIv1fTEJRFYKR |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-15 18:40 | gurjit@rxoom.com | 2026-06-15 | Aaihms Hospital | Surat |  | Dr. Hitesh Chitrode | Gurjit | CT-Spine | Spine Fixation |  |  | Yes |  | Partially | [5 empty] | NA | Pre operative CT scan used. Only fiducial registration was used for single level vertebra registration. Surgeon did two time registration in second registration got better registration accuracy. Surgeon confirmed the trajectory using navigation, used navigable awl, pedicle and compared with c arm on one pedicle, he said on navigation it is showing lateral on navigation. I suggest surgeon for accuracy, it can be improved with re-registration and but he didn’t go for it found navigation less value additive in spine and also mention reference moves with hammering. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.49 | 4.6 | 100 | https://drive.google.com/open?id=1C0p3fLZ5R6onObgK7oOPA2NlpMPx3i7J | https://drive.google.com/open?id=1Olei6nBclt79i37NDZjjxu0kFY6ITevF | https://drive.google.com/open?id=1aHoQd4EKM8srm8S3ivO1AAeMrmFvVeLv, https://drive.google.com/open?id=12yaRyLxK6u72xJmzdZ2wuI4eAXM2KHRZ, https://drive.google.com/open?id=10rZaPEflJ3XUYXDW3271v9oOMTc518TZ, https://drive.google.com/open?id=1nL3iMHTZkStOOkFldrmmQ2MvIrJl6aPc |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-16 20:14 | trushalgulhane1547@gmail.com | 2026-06-15 | Nizam hospital | Hyderabad |  | Dr. Akshay rajput and dr Vamsi Krishna | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Successful | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.16 | 2.22 | 0.85 | https://drive.google.com/open?id=1CeDCihPVWoKbmsX2SJoW6iwBAf_rFR80 | https://drive.google.com/open?id=1hcZKcl-YGNLjHx6rgYTAh6M4K1YLzPnQ | https://drive.google.com/open?id=1EClFf-xmTGss_mY41ZXrltXGFKvFMBxy |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-16 20:14 | trushalgulhane1547@gmail.com | 2026-06-16 | Medicover hospital securrendaba | Hyderabad |  | Dr. Shigakolli ramesh | Trushal Gulhane | Cranial | Awake craniotomy and tumor excision | Tumor excision |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of EM neuro navigation system | NA |  | Good | NA | Good | NA | NA |  | NA | NA | Trushal Gulhane | EM | V1.3.28 | Other: Please write Serial Number of the system | EM laptop system | NA | 0.02 | 4 | 0.67 | https://drive.google.com/open?id=1GwI_dpKbkASjK9DyY4SFfDxFjMHbiNUN | https://drive.google.com/open?id=1t96HT2qiZL0q5LWCHIuOTfxnVS1Nh4nL | https://drive.google.com/open?id=1T7SRcfU7UILhqkfIe1_Z97CoEdfyrcCf |  |  | https://drive.google.com/open?id=1TmJFAkol3WynzUOpSkK2Ap1lyk0xyyat |  | NO, this demo is not at an existing customer location.
2026-06-16 20:14 | hemarshitha1802@gmail.com | 2026-06-16 | Vydehi hospital | Bangalore | NA | Dr Vishwanath & Dr Ravi Teja | Hemarshitha H S | Cranial | CP Angle Tumor Excision |  |  | Yes |  | Completely |  |  | Scrolling of image was not happening in coronal view |  |  | NA | Preoperative CT data were loaded into the navigation system. A standard reference frame was used and registration was completed successfully. Registration accuracy was verified and accepted by the surgeons. The navigation system was used throughout the procedure. | NA | Scrolling of image was not happening in coronal view | Good | Good | Good | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.28 | NS OCR V01 A4 005 |  | NA | 0.26 | 3.4 | 1.7 | https://drive.google.com/open?id=1RQ-ju8shTlvP0f5hpM985WygxtTksGWH | https://drive.google.com/open?id=1VQYOZtnOsst5mxEQbFuVeA8C-LroXs5W | https://drive.google.com/open?id=1iX3q1TfHPdtjmljSEO9E5DZ95vkB8pHu |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-17 16:46 | gurjit@rxoom.com | 2026-06-17 | Aaihms Hospital | Surat |  | Dr. Maulik Patel | Gurjit | CT-Spine | C1-C2 Fixation |  | 1000 | Yes |  | Partially | [5 empty] | NA | Preoperative CT scan was loaded, spine clamp was attached on C2 and single level registration was done on C2 registration accuracy was good, surgeon was confirming trajectory on one side of pedicle which was showing good but to archive that trajectory he wants to use drill as navigable tool. Then surgeon did registration again for other side of pedicle still he wants to navigate drill. Dr find spine attachment spine clamp not ergonomic. | Spine Clamp design can be improved. Attachment for spine drill and other tools to make it navigable. |  | Disagree | Strongly Disagree | Agree | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.18 | 1.93 | 100 | https://drive.google.com/open?id=1aBw35TvOReEqyQ3wUj-cluTlFb-ETBCb | https://drive.google.com/open?id=19fD1EIAUB003KryfeYrq3IqqqkXhpdzD, https://drive.google.com/open?id=1Z7zfMsuRX5j2uMGYg8k9hp-dMmZLkMzx, https://drive.google.com/open?id=1_jKKKiS0tETnI9RjhHxkCk7Ydue9CuRE | https://drive.google.com/open?id=1VnfJEq-hizhSSuE9Nd3gx6Q9bWOBdw8V, https://drive.google.com/open?id=1Xd9oPJUlqrspxnFHahBuiM62GOeMqiQd, https://drive.google.com/open?id=1JMBwX0SI7DUFK-dT0Id1y8BOknoKHA6S, https://drive.google.com/open?id=18lbHagEgbXp6CCWw5oVDzsKz8oAb2wRU | https://drive.google.com/open?id=1Gx-4g8xPjGxd2_9j7Yn6V4DLU56KE8OV | https://drive.google.com/open?id=1f3HhfSFVTn27o3l1ajMUjwrjC8OVYip2 | https://drive.google.com/open?id=1LG2X-3DvjP7clDjjRXcMCowgUseBV7wt |  | NO, this demo is not at an existing customer location.
2026-06-18 13:06 | gurjit@rxoom.com | 2026-06-18 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit | Cranial | Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative MRI was loaded and standard reference was used, patient position was supine. Registration was done and accuracy was confirmed by surgeon. Navigation was used for its guidance throughout the surgery. | Surgeon wants to use suction as a navigable tool. |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.31 | 5.77 | 0.42 | https://drive.google.com/open?id=18_tqoCrS9Of1RZHIFABexeCQqsosqmZV | https://drive.google.com/open?id=1JzxVriLYJO4DVxgCvxUXMm4LCsSz4VCL | https://drive.google.com/open?id=1lRLUP1IeuY_iGC_cfNS1EewYSANUtOmN, https://drive.google.com/open?id=1Vyw5HpEJyNW7yxSyodPdrbPbu7gwfkCO |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-18 20:32 | gurjit@rxoom.com | 2026-06-18 | Avdhooth Hospital | Vadodara |  | Dr. Yogendra Bhutt | Gurjit | Cranial | VP Shunt |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded. Patient position was lateral, skull tracker was used. Registration and accuracy confirmed by surgeon. VP shunt guide tool was used for shunt placement. Navigation was used throughout the surgery for its guidance. | Tracker for different surgical tools. |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.11 | 6.02 | 0.52 | https://drive.google.com/open?id=1Z7zGOF3qWc32gBM5k5jxZQdCUqZ7wIho | https://drive.google.com/open?id=1IzWiXu14TtCqN7sDsik9U3BjBSwXH1by | https://drive.google.com/open?id=1TQDHlwXwGHQ6q-EAtrVR7D2ZRhlcJ94r, https://drive.google.com/open?id=1zb0N1R4ZAc30nHg52Wy6_sL4NlkopR4B |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-22 17:33 | gurjit@rxoom.com | 2026-06-22 | Avdooth Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit | Cranial | Tumor Excision |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded. Patient position was supine, skull tracker was used. Registration and accuracy confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | Probe tip should be more smooth. |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.11 | 5.57 | 0.55 | https://drive.google.com/open?id=1FdwrDbLe0nLPocKoJUKmc_LjNx5qNlND | https://drive.google.com/open?id=1VOdQQPvnbXeCUiWXvBSeZ9qJ3g_LRlT-, https://drive.google.com/open?id=1WkV1c4RTfmy90fpMhcBOd5KOGWzWlZhg | https://drive.google.com/open?id=1r2H03j4-gCnpzr2PLpCA65w0q6ggZxSP |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-23 15:08 | gurjit@rxoom.com | 2026-06-23 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit | Cranial | Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded. Patient position was supine, skull tracker was used. Registration and accuracy confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.31 | 5.08 | 0.34 | https://drive.google.com/open?id=1KFSBRzIpHwWTwRHoeRWY89ov56cpwitg | https://drive.google.com/open?id=1-EQAEhjgCHMEVX3AlTVVeGayLzBb9FAZ | https://drive.google.com/open?id=1DSp2gXD-pkWC9wbxxDFzxVOjuxjqg2l3, https://drive.google.com/open?id=1-9qKgJgQtGAnXGkGA0Rs4eIRoVdZ5-fS |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-27 13:10 | gurjit@rxoom.com | 2026-06-27 | Bhailal Amin General Hospital | Vadodara |  | Dr. Nimesh patel | Gurjit | CT-Spine | Lumber Fixation |  | 1000 | Yes |  | Partially | [5 empty] | NA | Pre operative CT scan used. Only fiducial registration was used for single level vertebra registration. Surgeon compared the accuracy with c arm on pedicle, he said on navigation it is showing mostly similar but lateral on navigation. I suggest surgeon for accuracy, it can be improved with re-registration and but he didn’t go for it. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 100 | 3.95 | 100 | https://drive.google.com/open?id=1_2VIO8mALLK4kgn0Z66nJmX5_JN6FQtE | https://drive.google.com/open?id=10-BPDX1D436LXoXQBZazY8rg4s7MeXJd | https://drive.google.com/open?id=1QsSranW9JNgtabqOAppM6eyzmA2spW2M, https://drive.google.com/open?id=12v3Pz3Fu3FfKkSJhP2DbbWOp2ptNOHLn, https://drive.google.com/open?id=17IqIn0TpZ0ZS5IhagfUSeuK7jvECXg_u, https://drive.google.com/open?id=1iqJckLjkQhZO29bs4WUPTxGzr9zEHdNh |  |  |  |  | NO, this demo is not at an existing customer location.
2026-06-30 11:10 | hemarshitha1802@gmail.com | 2026-06-25 | Jalappa hospital | Kolar | NA | Dr Somesh & Dr Madhushree | Hemarshitha H S | CT-Spine | C7 - D1 corpectomy with fixation |  | 25 | Yes |  | Completely |  | X ray reference was not detecting by the camera | X ray reference was not detecting by the camera and in the setup mode there was no option for Xray. |  | X ray module was not working in the system | NA | Initially surgeons planned for x ray based navigation but after acquired c arm shot of spine in lateral and AP, the x ray reference was not detecting by the camera so later they planned for CT based navigation. | To know exact reason why x ray based navigation was not working |  | Good | Good | Good | Need to use x-ray based navigation | NA |  | NA | NA | Faiz Khan | Optical | V1.2.28 | Other: Please write Serial Number of the system | NX010P003 | NA | 0.3 | 4.15 | 0.78 | https://drive.google.com/open?id=1m6-xBMemclF-HCtyFSdDrCp8LD9YVkXU | https://drive.google.com/open?id=1ksLV3tiiqq4_S4GhLOL0fks_-qwvBZw6 | https://drive.google.com/open?id=1UOB1U9_Imfy0Y2WTF1HopLHZp_h3cNOl |  |  |  |  | YES, this demo is at an existing customer location.
2026-06-30 20:02 | trushalgulhane1547@gmail.com | 2026-06-30 | Nizam hospital | Hyderabad |  | Dr. Prajwal and dr. Tirumal sir unit | Trushal Gulhane | Cranial | TNTS |  |  | Yes |  | Completely | [5 empty] | NA | Successful | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.16 | 2.43 | 0.76 | https://drive.google.com/open?id=1HwWMAxD9TXvki5NXWVkRtYsoENmTnZEj | https://drive.google.com/open?id=1NraWq3T5DjvglXJnQgyjUDxngq0hUU8k | https://drive.google.com/open?id=1IMqov6JNngz6_LDELLOtbSnSPbS8SPlE |  |  |  |  | YES, this demo is at an existing customer location.
2026-07-02 15:22 | trushalgulhane1547@gmail.com | 2026-07-01 | Nizam hospital | Hyderabad |  | Dr. Suchinda mam team | Trushal Gulhane | Cranial |  |  |  | Yes |  | Partially | [5 empty] | NA | Successful use for incision planning | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.2 | 2 | 1.35 | https://drive.google.com/open?id=1e-SmUAPf-_1HPQ4ryPYXuMinMAWssUUE | https://drive.google.com/open?id=1pUIl3IGVaV7iNzpmfOfG7iSMRRvGRA4O | https://drive.google.com/open?id=1MdWVByXWZg0Mae-WWJ-g4Vyq-f6o2REu |  |  |  |  | YES, this demo is at an existing customer location.
2026-07-18 15:30 | gurjit@rxoom.com | 2026-07-18 | SGVP Holistic hospital | Ahmedabad |  | Dr. Brijesh Panchal | Gurjit | Cranial | Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded. Patient position was supine, standard reference was used. Registration and accuracy confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | NA |  | Good accuracy | NA | Agree | NA | NA |  | One glyonce coating | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.07 | 6.41 | 0.65 | https://drive.google.com/open?id=1t7qSzg1mn08HkdXypGjN0tJfC2JM5Jph | https://drive.google.com/open?id=1eMXtXCLKQjJLO1VP-yH-cU647QGPfn1G, https://drive.google.com/open?id=1IVrn2Mio31ayP7LHgPkPOyxB_08leakj | https://drive.google.com/open?id=1PFF0jZdAAaIWKwf1MBNiEZhMKhFKrd35, https://drive.google.com/open?id=1esHEP79tvM_Yf_nW8DMObPH2yrngou5r |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-20 10:46 | trushalgulhane1547@gmail.com | 2026-07-16 | Noble hospital | Pune |  | Dr. Chandrashekhar raman | Trushal Gulhane | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | NA | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | V1.2.24 | Other: Please write Serial Number of the system |  | NA | 0.02 | 4 | 0.4 | https://drive.google.com/open?id=1mhYrkhVL3v1eOipsc-JCkylXcn2esoxX | https://drive.google.com/open?id=172en0P16WMZG2h3rZqjqTarvo-4PHSZY | https://drive.google.com/open?id=1gYcuswesjfbBtqucETIaje3nXMp6K6Gl |  |  |  |  | YES, this demo is at an existing customer location.
2026-07-20 23:51 | gurjit@rxoom.com | 2026-07-20 | SGVP Holistic hospital | Ahmedabad |  | Dr. Brijesh Panchal | Gurjit | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded. Patient position was lateral, standard reference was used. Registration was done and accuracy was confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Delhi Distributor | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.14 | 4.23 | 0.97 | https://drive.google.com/open?id=1hrD8Mkjkb8nKdoyq6MPz5FybKVTXqxOr | https://drive.google.com/open?id=1DhsGhFIBqw15zZ1liIsaRjihBiJA8zAj | https://drive.google.com/open?id=1Bct3AXUyD89BEhYWifgUox4weVwPiMe4 |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-23 14:32 | gurjit@rxoom.com | 2026-07-23 | BMHRC | Bhopal |  | Dr. Sandeep Sorte | Gurjit | Biopsy, Any (Cranial, Full body) |  |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded. Patient position was Supine, standard reference was used. Registration was done and accuracy was confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP004 | NA | 100 | 2.7 | 0.28 | https://drive.google.com/open?id=1xS77a6UmZuqETTUEdK1eE1DVB-AKnn3U | https://drive.google.com/open?id=1i_IV-6AX8SrGIIvzNXpK7mBC0qBSZ6I8 | https://drive.google.com/open?id=1IG6vwOU3VVccmCJi4IT5IFdXERxAaq4G, https://drive.google.com/open?id=185RTQv1Z0eUFCTjRHtiMJOP2bSS0_l9Y, https://drive.google.com/open?id=1dbCN6tUQeIfCFkA3nd3fsq_39-ZH5VSn |  |  |  |  | YES, this demo is at an existing customer location.
2026-07-23 14:37 | hemarshitha@rxoom.com | 2026-07-22 | First Neuro Hospital | Mangalore | NA | Dr Skanda Murthy | Hemarshitha H S | Cranial | VP shunt |  |  | Yes |  | Completely |  |  | Following the registration process, the 3D model was not visible on the system display. |  |  | NA | Preoperative CT data was imported into the navigation system. A skull tracker was applied and patient registration was successfully completed. Registration accuracy was verified and accepted by the operating surgeon. The navigation system was utilized throughout the entire procedure. | NA |  | Good | Good | Good | NANA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.24 | NS OCR V01 A4 005 |  | NA | 0.16 | 5.72 | 0.72 | https://drive.google.com/open?id=1AFAAkqJO_-JTTQZqiEAkpBbHQ-TgVXYG | https://drive.google.com/open?id=1xxU3WgR_knfmW6hTjLAje00qoW1aegYV | https://drive.google.com/open?id=1aItVGoCquMpoR2oEB4wVz45TXwHLBT1c |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-23 15:01 | hemarshitha@rxoom.com | 2026-07-22 | First Neuro Hospital | Mangalore | NA | Dr Skanda Murthy | Hemarshitha H S | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI data was imported into the system. A standard reference frame was applied and patient registration was successfully completed. Registration accuracy was verified and accepted by the surgeon. The system was utilized throughout the entire procedure. | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.24 | NS OCR V01 A4 005 |  | NA | 0.18 | 4.42 | 0.46 | https://drive.google.com/open?id=1ppiFpv42V0vXvBZgQJ8Q9Gg2lhHT6NOo | https://drive.google.com/open?id=1m2qVHdOhhySrZFsuY1EhRCIGVnfY858C | https://drive.google.com/open?id=1kaeguk545kOQeIeWhMnlMDTe-hOGIfHA |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-23 15:23 | hemarshitha@rxoom.com | 2026-07-21 | First Neuro Hospital | Mangalore | NA | Dr Skanda Murthy | Hemarshitha H S | Cranial | Right Periventricular Tumor Excision |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI data was imported into the navigation system. A standard reference frame was applied and patient registration was successfully completed. Registration accuracy was verified and accepted by the operating surgeon. The optical navigation system was utilized throughout the entire procedure to ensure precise guidance. | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.24 | NS OCR V01 A4 005 |  | NA | 0.24 | 5.42 | 0.93 | https://drive.google.com/open?id=1AfhHPfAFcAu4SCKYNaBqSlmOwoyjNjbQ | https://drive.google.com/open?id=1-p0xYs8oYV2MzkIXnF6lfXGWf9ciE-uL | https://drive.google.com/open?id=1bhWF6Ug7KkdcZHvVzMbEUJMFfjfNjfkR |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-25 21:18 | hemarshitha@rxoom.com | 2026-07-15 | Vydehi hospital | Bengaluru | NA | Dr Vishwanath | Hemarshitha H S | CT-Spine | L3-L5 screw fixation |  | 35 | Yes |  | Partially | While registration, the surgeon collected surface points only on the exposed vertebral surface, as the surrounding muscles limited access and prevented collection of additional registration points. He felt that the its accurate at the vertebral surface level and therefore continued using the system until planning the trajectory for deeper instrumentation. |  |  |  |  | NA | While registration, the surgeon collected surface points only on the exposed vertebral surface, as the surrounding muscles limited access and prevented collection of additional registration points. He felt that the its accurate at the vertebral surface level and therefore continued using the system until planning the trajectory for deeper instrumentation. | NA |  | Good at the surface level | NA | NA | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.24 | NS OCR V01 A4 005 |  | NA | 0.36 | 1.55 | 1.03 | https://drive.google.com/open?id=1mdPNEG2GflyX5yZNOJa6ZszHnqags2Yj | https://drive.google.com/open?id=1emp2_SDNRI9HHQ3F0lg9_GMoKGNycS07 | https://drive.google.com/open?id=1nG8BMJvdFly6tGQ4P2qAEyMUXrkGdda5 |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-25 21:31 | hemarshitha@rxoom.com | 2026-07-14 | Vydehi Hospital | Bangalore |  | Dr Benak | Hemarshitha H S | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT data was loaded into the navigation system. A standard reference frame was applied and patient registration was performed by Dr. Ravi. Registration accuracy was verified and accepted by the surgeon. The navigation system was utilized throughout the entire procedure. | NA |  | Good | Good | Good | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.24 | NS OCR V01 A4 005 |  | NA | 0.16 | 3.5 | 0.61 | https://drive.google.com/open?id=1kJRHf--VItYSZTXGRm8tB8CU_rNPCZeu | https://drive.google.com/open?id=1mzdSs9YtGGYYc33M9qOVpdwXi0_j239l | https://drive.google.com/open?id=1Nq_ImBtpC19TQ9K3Y9m3M3x6xp4WxyDp |  |  |  |  | NO, this demo is not at an existing customer location.
2026-07-30 19:44 | gurjit@rxoom.com | 2026-07-29 | Sri H N Reliance Foundation | Mumbai |  | Dr. Rajan shah and Dr. Uttaresh | Gurjit Singh | Cranial | Tumour removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI scan was loaded and Tractography data was used for fiber tracking and surgery planing. Patient position was Supine, standard reference was used. Registration was done and accuracy was confirmed by surgeon. Navigation was used throughout the surgery for its guidance. | Have a live tracking biopsy needle and microscope integration. |  | Agree | Agree | Strongly Agree | Surgeon compared striker navigation fiducial registration with our navigation and said in strike system after three fiducial point capturing it start the mapping on it helps to capture the rest of the fiducial point accurately. | NA |  | NA | NA | Gurjit Singh | Optical | 1.3.28 | NS OCR V01 A4 006 |  | NA | 0.21 | 4.26 | 0.69 | https://drive.google.com/open?id=1xWMkyGFJGwzL59alOWTrn7TKU534wEPl | https://drive.google.com/open?id=1QG20U5V5IOkN4--OU0kBKCt4QjgFZtFz | https://drive.google.com/open?id=1xUGS6JA5K9YX5HT4KeYof4FlPoJ7e0gZ, https://drive.google.com/open?id=1N4zPRuv5DolB7jGfFC8-iLXnLEGojlsD | https://drive.google.com/open?id=1z7fKO3rU_FkY0iObFQAOrQf_tLww0WBs |  | https://drive.google.com/open?id=1WCv7K-qXeV78dyvXdlUuVA21fuTQGzHs |  | NO, this demo is not at an existing customer location.
2026-08-01 11:54 | hemarshitha@rxoom.com | 2026-07-27 | First Neuro hospital | Mangalore | NA | Dr Skanda Moorthy | Hemarshitha H S | Cranial | VP Shunt |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT data was imported into the navigation system. A skull tracker was applied and patient registration was successfully completed. Registration accuracy was verified and accepted by the operating surgeon. The navigation system was utilized throughout the entire procedure. | NA |  | Strongly Agree | Strongly Agree | Strongly Agree | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.24 | NS OCR V01 A4 005 |  | NA | 0.18 | 4.8 | 0.63 | https://drive.google.com/open?id=1qyU5B4_D_mXwrjG0_Qjns_az1OjcWTpg | https://drive.google.com/open?id=1r8neUS4h0fepWVnotZVXMu6J6uA5z6uN | https://drive.google.com/open?id=1djXAYKbM1W8-FVB061ZlQm5ERVBzsxE6 |  |  |  |  | NO, this demo is not at an existing customer location.
2026-08-01 12:18 | hemarshitha@rxoom.com | 2026-07-13 | Vydehi hospital | Bangalore | NA | Dr Vishwanath | Hemarshitha H S | Cranial |  |  |  | Yes |  | Completely |  | No | No | During data import, CT Angiogram and MRI datasets were received from the hospital in DICOM format. - The CT dataset imported successfully into our navigation system. - The MRI dataset did not import. The same MRI dataset was verified to load correctly in the Stryker navigation system. - After discussion with Anitha ma'am and the Radiology HOD, it was identified that the DICOM format provided was not fully compatible with our system. - A revised dataset in a compatible DICOM format was obtained from Radiology and successfully loaded into our system. Thereafter, the procedure proceeded with navigation. |  | NA | Preoperative MRI data was loaded into the navigation system. A standard reference frame was applied and patient registration was performed by Dr. Vishwanath. Registration accuracy was verified and accepted by the surgeon. The navigation system was utilized throughout the entire procedure. | NA |  | Strongly Agree | NA | Agree | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.24 | NS OCR V01 A4 005 |  | NA | 0.2 | 10.47 | 1.54 | https://drive.google.com/open?id=1NBr7nFq59USJ4LwubliPr8qcfOeuw2d2 | https://drive.google.com/open?id=1j7Ptqizl0zj2VjLSp6jotx_m7PgHBxlH | https://drive.google.com/open?id=19_NkCbh2YwIqcA9d-GcHu0gmUctEBXg9 |  |  |  |  | NO, this demo is not at an existing customer location.
2026-08-01 12:34 | hemarshitha@rxoom.com | 2026-07-11 | Vydehi hospital | Bangalore | NA | Dr Benak | Hemarshitha H S | Cranial | Right Parietal Craniotomy and tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT data were loaded into the navigation system. A standard reference frame was used and registration was completed by Dr Benak successfully. Registration accuracy was verified and accepted by the surgeon. The navigation system was used throughout the procedure. | NA |  | Agree | Agree | Agree | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.24 | NS OCR V01 A4 005 |  | NA | 0.24 | 6.46 | 0.59 | https://drive.google.com/open?id=1NlRGwYN105lRFy9dLNy94g04hBxk_wYD | https://drive.google.com/open?id=1SJNkEwQ-lW1kME6y44FjxEQNzN6fR4ih | https://drive.google.com/open?id=1cLSt5cN6QQ9NATakVmKmW21-M1eUzbAR |  |  |  |  | NO, this demo is not at an existing customer location.
2026-08-01 12:51 | hemarshitha@rxoom.com | 2026-07-10 | Vydehi hospital | Bangalore | NA | Dr Ananth & Dr Vishwanath | Hemarshitha H S | X-Ray Spine |  |  | 50 | No | Intraoperative AP and lateral fluoroscopic images were acquired using the C-arm. The images were transferred to navigation system. However, due to the patient's obese body, an adequate lateral fluoroscopic image could not be obtained, making accurate navigation registration difficult. As navigation accuracy could not be ensured, the surgeons decided to proceed with the surgery without navigation assistance. | Not used |  |  | During image processing, it was observed that the C-arm was unable to export images in DICOM format and only supported JPG image export. Initially, the navigation system did not support direct conversion of JPG images into DICOM format. After discussing the issue with Anitha ma'am, it was suggested to first convert the JPG images into PNG format and then convert the PNG images into DICOM format. | - The C-arm supported image export only in JPG format and did not provide DICOM images. - Image conversion from JPG to DICOM was achieved by converting the JPG files to PNG format first and then generating DICOM images. |  | NA | Intraoperative AP and lateral fluoroscopic images were acquired using the C-arm. The images were transferred to navigation system. However, due to the patient's obese body, an adequate lateral fluoroscopic image could not be obtained, making accurate navigation registration difficult. As navigation accuracy could not be ensured, the surgeons decided to proceed with the surgery without navigation assistance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | 1.2.24 | NS OCR V01 A4 005 |  | NA | 00 | 00 | 00 | https://drive.google.com/open?id=1ge1to-Hi_KhfnEHxqyYwTXtgHOX63aEo | https://drive.google.com/open?id=1NnXEqPvUgO3Z8qNd9YXERnYiPwboTIqe | https://drive.google.com/open?id=1XFVBFcJ9OmroOyD1aGxh-u2hhPbsvIzF |  |  |  |  | NO, this demo is not at an existing customer location.
2026-08-10 14:24 | gurjit@rxoom.com | 2026-08-09 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | C1-C2 plating |  | 10 | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was loaded and crop functionary was used to crop to keep required area of CT scan. Clap was attached on C2. One at a time single level vertebra was registered for C1 and C2 and surgeon. Surgeon used the navigation for its guidance. | Requested better spine clamp |  | NA | Need improvement in spine attachment | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 0.09 | 3.64 | 0.32 | https://drive.google.com/open?id=1OSEz8yWmkaZZrSmOi-8NgW7q5opeKZVk | https://drive.google.com/open?id=1uhIc5ZNeTbXuEJvoGoWQNQb_Y2ccGhCL | https://drive.google.com/open?id=1IPqpa-M9rTXltf56JwUl_Kh6SHpGBkCa |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-12 20:07 | gurjit@rxoom.com | 2026-08-11 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | D11-L1 fixation |  | 1000 | Yes |  | Partially | [5 empty] | NA | Preoperative CT scan was used and to navigate on D12 surgeon used the skull tracke, place half part on D11 spinous process with one side two screws and for D11 used the spine clamp as a reference and used Pedicle finder and used the navigation for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | While using Pedicle finder, pedicle finder tracking and spine clamp reference was clashing. Spine clamp reference acting as tool and pedicle finder tool as a reference, which result pedicle finder was showing way off to spine. | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 0.2 | 5.34 | 100 | https://drive.google.com/open?id=1QOk_MWSs-D_NZ9mqJL_5SPf9hHu9oQtX | https://drive.google.com/open?id=1Zx-0mXgLc8GBdijSeOFwNI4VMnhRrNAL, https://drive.google.com/open?id=1su71chxHBI_-WdgRdbh5uwv7SlbApE0w, https://drive.google.com/open?id=1pstVsxFBiezgXZlBnNXf3RN_UUIlmmNS | https://drive.google.com/open?id=1977fUz1axjWuGj2ujjOZqtnfCgb70NzM, https://drive.google.com/open?id=1OuL_OPjB8vIo9y7BG-lf-ImvPRMeQqmO |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-12 20:31 | gurjit@rxoom.com | 2026-08-11 | Avdhoot Hospital | Vadodara |  | Yogendra Bhatt | Gurjit Singh | CT-Spine | C5-C7 fixation |  | 1000 | Yes |  | Partially | [5 empty] | NA | Pre operative CT scan was loaded and CT spine clamp was used and single level registration was done on C5 and navigation was used for its guidance after the camp was moved they stopped using the navigation. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 0.21 | 3.85 | 100 | https://drive.google.com/open?id=1rT7hhU7DVGwAbQRHPW3qnvHIHQ4hQHQJ | https://drive.google.com/open?id=1JTJAbVQTnpHnclI7aLPwEJof9vO1CaFS | https://drive.google.com/open?id=1WfjKQO3IOJMcNqSmIE4-GYd6p25JjkLh |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-15 11:21 | gurjit@rxoom.com | 2026-08-14 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | L4-L5 Facet Fixation |  | 20 | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was loaded and skull tracker with stitches was used and single level registration was done on L5 and navigation was used for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 0.18 | 7.12 | 0.2 | https://drive.google.com/open?id=1OZj7Kv2G0mxVxxRGxdbr_9GcQlUv8h5V | https://drive.google.com/open?id=1OeMn7SrLgEAUNs1uyU4JfTdC94WyyEtH | https://drive.google.com/open?id=1tap3DDSccWt18ZjNFjpC3FdOiWqIjM4d |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-15 11:46 | gurjit@rxoom.com | 2026-08-14 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | C1-C2 Fixation |  | 20 | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was loaded and skull tracker with stitches was used and single level registration was done on C2 and VP shunt guide tool was used as a probe and navigation was used for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 100 | 3.14 | 100 | https://drive.google.com/open?id=1rhTcIitNoaaJoBH4uq-T8ZvkviXLQ00m | https://drive.google.com/open?id=1kOwworvIfGS913U5gEanqKJwu4Unjssc, https://drive.google.com/open?id=1omyODOroiICx-zu94d3VQAEXPmhYj5Ev | https://drive.google.com/open?id=110c35b6OZAd3h8v1ZNhZUqENFjODO6bE |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-17 11:15 | hemarshitha@rxoom.com | 2026-08-04 | Jalappa hospital | Kolar | NA | Dr Somesh | Hemarshitha H S | Cranial |  |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI data was imported into the navigation system. A skull tracker was applied and patient registration was successfully completed. Registration accuracy was verified and accepted by the operating surgeon. The navigation system was utilized throughout the entire procedure. | NA |  | Good | NA | NA | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V 1.2.24 | Other: Please write Serial Number of the system | NX010P003 | NA | 0.24 | 3.85 | 0.93 | https://drive.google.com/open?id=1uvfHs1cJ5UrvfHPxDw0rhIlEnjSQaMd9 | https://drive.google.com/open?id=1BzyvpnOYmDFjl_sJXh0ALjQ2q96OIeMw | https://drive.google.com/open?id=1sG_xOsfA97wV7TC37DKgqXY-oRi5Z1AN |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-17 12:09 | hemarshitha@rxoom.com | 2026-08-11 | Vydehi hospital | Bangalore | NA | Dr Benak | Hemarshitha H S | Cranial | Frontoparietal Craniotomy and tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI data were loaded into the navigation system. A standard reference frame was used and registration was completed successfully. Registration accuracy was verified and accepted by the surgeon. The navigation system was used throughout the procedure. | NA |  | Good | NA | NA | NA | NA |  | NA | NA | Hemarshitha H. S. | Optical | V1.2.24 | NS OCR V01 A4 005 |  | NA | 0.16 | 2.96 | 1 | https://drive.google.com/open?id=1aEkZEV4b52Rq2VZ-dxo0y9QaQbzS5MVQ | https://drive.google.com/open?id=1zBuzucYINsD9FuEffmE4T0BnQ2WxP3rc | https://drive.google.com/open?id=1_JgQFmLZC-jxMPfwEZgrHxCT_ODd5f_u |  |  |  |  | NO, this demo is not at an existing customer location.
2026-08-21 20:47 | gurjit@rxoom.com | 2026-08-21 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | L4-L5 Facet fixcation |  | 20 | Yes |  | Completely | [5 empty] | Skull tracker with stitches was used in for spine surgery VP shunt guide tool was used as a probe | Pre operative CT scan was loaded and skull tracker with stitches was used and single level registration was done on L5 and VP shunt guide tool was used as a probe and navigation was used for its guidance. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.28 | Other: Please write Serial Number of the system | NX01OP006 | NA | 100 | 3.6 | 0.45 | https://drive.google.com/open?id=16KbTUDOCUwPckt2APIr8Bwz8ZJm8iojq | https://drive.google.com/open?id=1GWlZmhFFNWIzI_wztFMXiV1d4ThejRhD | https://drive.google.com/open?id=1KvdzjsELEuVRajWsSWYGW7Lukbn63ox4, https://drive.google.com/open?id=1v-WAAxymh_GOhxdzEp5BjZ0mQYdx9Bss |  |  |  |  | YES, this demo is at an existing customer location.
2026-08-28 15:51 | trushalgulhane1547@gmail.com | 2026-08-27 | H.N. Reliance Hospital | Mumbai | NA | Dr. Rajan and Dr. Uttareshwar | Trushal Gulhane | Cranial | Prone position tumor excision |  |  | Yes |  | Completely | [5 empty] | NA | Successful demonstration of neuro navigation system | Integration with microscope |  | Very good | NA | Good | NA | NA |  | NA | NA | Mumbai Distributor | Optical | V1.3.28 | Other: Please write Serial Number of the system |  | NA | 0.2 | 6.23 | 2 | https://drive.google.com/open?id=1-s6AXVZwk3ZirSM0iKStn4YMCqT8-s5R | https://drive.google.com/open?id=1nc9nYDAG_0G5pqLPiSXEeC3aQ7F5Nm6p | https://drive.google.com/open?id=1ikpjY2mMtVDNE2P54hgCVNF9Tgxrd7e0 |  |  | https://drive.google.com/open?id=18-Wxh8_mympH8apexS2DgnBims0OiTfG |  | NO, this demo is not at an existing customer location.
2026-09-08 19:28 | gurjit@rxoom.com | 2026-09-08 | Darsh Hospital | Kalaburagi |  | Dr. Shashank Ramdurgh | Gurjit Singh | Biopsy, Any (Cranial, Full body) | Brain biopsy |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative CT scan was loaded. Registration was done and accuracy was confirmed by surgeon.Surgeon planned the two trajectory. Biopsy region was more than 3 cm and depth was within range of 5cm. Surgeon took approx 8 biopsy samples 4 on each trajectory with variation of more depth by 4-8mm | NA |  | NA | NA | NA | Trajectory alignment support Lyla arm can be improved | NA |  | NA | NA | Faiz Khan | Optical | 1.2 | Other: Please write Serial Number of the system | NSOCRV01A3002 | NA | 0.52 | 5.23 | 0.28 | https://drive.google.com/open?id=1jmr7F9Z3FyaHYeDs_9PQqiSRHybpfOxN | https://drive.google.com/open?id=10B704wDs_bDQcmAr5Z9Lt3hxgsnAFBgv | https://drive.google.com/open?id=1jwKYqswsKn8Npm1kiEUB5kzALYq691ir, https://drive.google.com/open?id=1xdBha_6PSJbFqHa9FxlgE6HriJl5Df76 |  |  |  |  | YES, this demo is at an existing customer location.
2026-09-25 09:41 | gurjit@rxoom.com | 2026-09-24 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | Cranial | Brain Aneurism |  |  | Yes |  | Completely | [5 empty] | NA | Pre operative CT ANGIO scan was loaded and patient position was supine. Standard reference was used, Patient registration was done and accuracy was confirmed by surgeon new bayonet forcep tool attached with tracker was used. | NA |  | NA | NA | NA | NA | NA |  | NA | NA | Faiz Khan | Optical | 1.3.30 | Other: Please write Serial Number of the system | NS TRL 02 MF 211 | NA | 100 | 3.5 | 0.46 | https://drive.google.com/open?id=1AzZ5T8poW5uY2I_V6yJ4-BY0QBbtndmC | https://drive.google.com/open?id=1DhJGA8TvIWIGsM55ddFuBH0SR7FUsPrE | https://drive.google.com/open?id=1w7m2L9WaS4gK_-gZnOUqDqJANkb6Zb25 |  |  |  |  | YES, this demo is at an existing customer location.
2026-09-26 21:10 | gurjit@rxoom.com | 2026-09-26 | Avdhoot Hospital | Vadodara |  | Dr. Yogendra Bhatt | Gurjit Singh | CT-Spine | L2-L4 Fixation |  | 30 | Yes |  | Completely | [5 empty] | NA | Pre operative CT scan was loaded, cranial module was opted to use skull tracker for reference, block with mini plane was applied on patient back using stitches and cranial self tapping screws. Single level registration was done on L2, bayonet tool, navigation guidance was used for pedicle finding and screw placement. | NA |  | NA | NA | NA | Bigger cranioplasty screws can be used to apply skull tracker on spinous process through skin. | NA |  | NA | NA | Faiz Khan | Optical | 1.3.30 | Other: Please write Serial Number of the system | NS TRL 02 MF 211 | NA | 0.11 | 6.48 | 100 | https://drive.google.com/open?id=1TleqUGiU3kfSTjRCRFnFeeL87X86W_Lq | https://drive.google.com/open?id=1LU76Jhn3IptGqp0F7H7S6-jPNeVW3dT1 | https://drive.google.com/open?id=18vIeuf9pKd0Wc6CskIjSUfrysXdgXLHZ, https://drive.google.com/open?id=1rMdpnRvEzY_4g-lOvZsPgMBGAXORpJyn |  |  |  |  | YES, this demo is at an existing customer location.
2026-09-30 12:48 | gurjit@rxoom.com | 2026-09-29 | Avdhoot Hospital | Vadodara |  | Yogendra Bhatt | Gurjit Singh | Cranial | Tumor removal |  |  | Yes |  | Completely | [5 empty] | NA | Preoperative MRI was loaded, patient position was lateral. Registration was done and accuracy was confirmed by surgeon, craniotomy was marked using navigation was used for its guidance. | NA |  | NA | NA | NA | Develop microscope integration, HRS is providing integration in some of the company microscope. | NA |  | NA | NA | Faiz Khan | Optical | 1.3.30 | Other: Please write Serial Number of the system | NS TRL 02 MF 211 | NA | 100 | 7.59 | 0.55 | https://drive.google.com/open?id=1_az1j-5bjHlXs1gDutUxvh3l2GyMlKku | https://drive.google.com/open?id=1i4g3U6WUGb2IpU3XCt-v-NQzjxmDaKWN | https://drive.google.com/open?id=1pBOtIA13YSJJblUR8TnJxqACMxuyLl3m |  |  |  |  | YES, this demo is at an existing customer location.